## NCIT Table
Patient Selection (L:Low; U: Unknown; H: High | Index Test | Reference Standard | Flow and Timing | Overall Risk of Bias | Applicability Concerns | Year | Author | Scenario | Tool brand | Measurement Location | Measurement distance | Group size | Age(years) | Male | Female | Ambient temperature(°C) | Reference results distribution(°C) | Reference temperature | Reference tool | Correlation coefficient | Mean Absolute Error(MAE)(°C) | Mean Difference(MD) /Bias(°C) | standard deviation(SD) (°C) | Reference cut-off temperature (°C) | NCIT cut-off temperature | sensitivity | specificity | Area under curve(AUC) | positive prediction value (PPV) | negative predictive value(NPV)
L | L | L | L | L | L | 2012 | Teran, C. G. [1] | hospital | Thermofocus, model 01500, TECNIMED, Varese, Italy | forehead |  | 434 | 14.6 ± 10.7 months | 208 | 226 | 24 to 28 | 37.9 ± 0.9 | rectal | rectal glass mercury thermometer | 0.952 |  | 0.029 | 0.01 | 38 |  | 97% | 97% |  | 95.2% | 98.1%
H | L | L | L | H | H | 2013 | Basak, Tulay [2] | room | Thermosense | forehead | 5cm | 452 | 19.66 ± 0.94 | 203 | 249 |  | 36.22 ± 0.10 | oral | chemical dot thermometer | 0.328 |  | -0.49 | 0.41
[18 empty] | tympanic | tympanic thermometer | 0.805 |  | -0.57 | 0.35
U | L | L | L | U | U | 2011 | Chiappini, Elena [3] | hospital | Thermofocus | forehead |  | 251 | range 3.0 to 8.6 | 127 | 124 | 24.15 ± 1.81 | 37.18 ± 0.96 | axillary | mercury-in-glass thermometer | 0.914 |  | 0.07 | 0.35 | 38 |  | 88.7% | 89.9% |  | 70% | 97%
U | L | L | L | U | U | 2008 | Williams, E. M. [4] | room | Raytek, Raynger MX | facial skin |  | 169 | 25 ± 8 (range 18 to 54) |  |  | 20 to 22 | 36.7 ± 0.5 | tympanic | tympanic thermometer |  |  | -2.5 to 2.7°C
U | U | H | U | H | H | 2020 | Costanzo, Sandra [5] | room |  | forehead |  | individual |  |  |  | 29.17 ± 3.04 (range 24 to 34) | 36.47± 0.202 |  |  |  | 0.173 | -0.1667 | 0.1647
H | L | L | L | H | H | 2022 | Kameda, Norihiro [6] | hospital | TM-1621; A&D Medical | forehead |  | 65 | 66.8 ± 6.4 (range 50 to 80) | 65 | 0 | 24 | 36.6 ± 0.1 | skin (core temperature) | zero-heat-flux cutaneous thermometer | 0.8 |  | -0.18 | 0.22
L | L | L | L | L | L | 2021 | Erdem, Nurettin [7] | hospital |  | forehead |  | 151 | 7.95±4.48 (range 2 to 18) | 81 | 70 | 24 | 37.81±1.2 | axillary | mercury-in-glass thermometer |  |  |  |  | 38 |  | 69% | 100% |  | 100% | 93%
L | L | L | L | L | L | 2017 | Hakan, Nilay [8] | hospital | ThermoFlash® LX-26 | temporal | 10cm | 927 | 10.73±16.32 (range 0.01 to 82) | 358 | 569 |  | 37.1±0.5 | rectal | rectal thermometer |  |  | -0.03 | 0.45 | 38 |  | 48.5% | 99.5% |  | 88.9% | 96.3%
[17 empty] | 37.4±0.9 | tympanic | tympanic thermometer |  |  | 0.14 | 0.596
[17 empty] | 36.5±0.7 | axillary | axillary thermometer |  |  | 0.59 | 0.49
[17 empty] | 36.4±0.5 | oral | oral thermometer |  |  | 0.33 | 0.469
L | L | L | L | L | L | 2018 | Franconi, Ilaria [9] | hospital | Hartmann Thermoval Duo Scan (Model 925082; Hartmann, Germany) | forehead |  | 205 | 4.89 ± 3.86 (range 0 to 14) | 110 | 95 |  | 37.52±1.09 | axillary | digital axillary thermometer |  |  | -0.41 | 0.81
L | L | L | L | L | L | 2020 | Huang, Fanyu [10] | hospital | HYLOGY MD-H26 | forehead | 5cm | 26 | 48.6 ± 17.5 | 14 | 12 | 25 | 37.84±1.02 | axillary | mercury-free thermometer | 0.47 |  | −1.1892 | 0.9049
L | L | L | L | L | L | 2022 | Lai, Fan [11] | hospital | JXB-178 | neck | 5cm | 1860 in total, 1304 children | 28.56 ± 7.25(adult),3.45 ± 2.85(children) | 683 | 1177 | 22.5±4.09 (range 9 to 32) | 37.10±0.92 | axillary | mercury axillary thermometer |  |  |  |  | 37.3 | 36.65 | 95.6% | 62.3% | 0.922
[10 empty] | temporal | [14 empty] | 36.45 | 96.1% | 61.7% | 0.911
[10 empty] | forehead | [14 empty] | 36.35 | 96.8% | 39.8% | 0.893
[10 empty] | wrist | [14 empty] | 36.15 | 95% | 34.3% | 0.824
L | L | L | L | L | L | 2021 | Pana, Bogdan C. [12] | hospital |  | skin | 1 to 5cm | 5231 | 53.97±17.95(range 0.5 to 91) | 2772 | 2459 | [9 empty] | 37.3 | 9.43% | 99.15%
L | L | L | L | L | L | 2022 | Lai, Fan [13] | hospital | JXB-178 | neck | 5cm | 904 in total, 669 children | 29.63±6.30(adult),3.71±3.25(children) | 390 | 514 | 19.8±3.72 (range 14 to 29) | 37.33±0.97 | axillary | mercury axillary thermometer |  |  |  |  | 37.3 | 37.35 | 86.6% | 84.6% | 0.92
[10 empty] | temporal | [14 empty] | 36.65 | 81.8% | 85.2% | 0.912
[10 empty] | forehead | [14 empty] | 36.65 | 78.8% | 83.6% | 0.89
[10 empty] | wrist | [14 empty] | 36.75 | 62.8% | 87% | 0.827
L | L | L | L | L | L | 2002 | Kocoglu, H [14] | hospital | Braun ThermoScan IRT | tympanic |  | 110 | 7.7±2.2(range 5 to 10) | 66 | 44 |  | 38.18±1 | rectal | mercury thermometer |  |  | 0.55 | 0.4
H | L | L | L | H | H | 2019 | Moran-Navarro, Ricardo [15] | room, rest | Medisana, Germany | forehead | 5cm | 12 | 25±3.9 | 12 | 0 | 20.1 ± 0.3 | 36.7±0.8 | core(gastrointestinal) | ingestible sensor |  |  | -0.1 | 0.89
[8 empty] | room, exercise, no sweat,no wind | [8 empty] | 38.3±0.9 |  |  |  |  | 0.1 | 1.22
[8 empty] | room, exercise, no sweat,wind | [8 empty] | 38.3±0.9 |  |  |  |  | -1.4 | 1.09
[8 empty] | room, exercise, sweat,no wind | [8 empty] | 38.3±0.9 |  |  |  |  | 0.1 | 1.147
[8 empty] | room, exercise, sweat,wind | [8 empty] | 38.3±0.9 |  |  |  |  | -1.9 | 1.045
[8 empty] | room, rest | Tecnimed, Italy | forehead | 6cm | [5 empty] | 36.7±0.8 |  |  |  |  | 0.3 | 0.95
[8 empty] | room, exercise, no sweat,no wind | [8 empty] | 38.3±0.9 |  |  |  |  | 0.2 | 1.02
[8 empty] | room, exercise, no sweat,wind | [8 empty] | 38.3±0.9 |  |  |  |  | -1.2 | 0.99
[8 empty] | room, exercise, sweat,no wind | [8 empty] | 38.3±0.9 |  |  |  |  | 0.3 | 0.99
[8 empty] | room, exercise, sweat,wind | [8 empty] | 38.3±0.9 |  |  |  |  | -1.6 | 1.02
L | L | L | L | L | L | 2022 | Kim, Yeon-M [16] | hospital | Hubdic Thermofinder S2, HFS-710, Gyeonggi-do, Korea | forehead | 2 to 3 cm | 255 | 9.05 ± 5.39 | 140 | 115 | 25 to 27 | 37.39 ± 0.66 | tympanic | infrared tympanic thermometer | 0.87 |  | -0.15 | 0.295 | 38 |  | 81% | 98% | 0.95 | 90% | 96%
U | H | L | U | H | H | 2017 | Ray, Partha Pratim [17] | room | MLX90614 | forehead | 1.5 to 30 cm | 37 | 38± 10.89 | 18 | 19 |  | 36.3±0.056 | forehead(core) | CoreTemp CM-210, core temperature sensor |  |  | -0.7
H | L | L | U | H | H | 2022 | Ajcevic, Milos [18] | hospital | HG01，Comelit | forehead | 10cm | 30 | 71.9 ± 18.9(range 22 to 97) | 24 | 6 |  |  | tympanic | infrared tympanic thermometer, TEGENIUS3，Tyco Healthcare | 0.74 |  | -0.14 | 0.34
H | U | L | U | H | U | 2021 | Goh, Nicholas Wei-Jie [19] | room | MLX90614 | forehead | 2 to 4 cm | 5 | >21 |  |  | 20 to 22 | 36.34±0.375 | oral | digital thermometer,MC-343F |  |  |  | 0.3
L | L | L | L | L | L | 2021 | Rinanto, Noorman [20] | room | MX90614 | forehead | 2 to 30 cm | 10 | [5 empty] | axillary | MT 16C2 digital thermometer |  | 0.1 |  | 0.03
[9 empty] | TD-1242 | [11 empty] | 0.48 |  | 0.17
L | L | L | L | L | L | 2022 | Mah, Aaron James [21] | room | SCT01 | forehead | 2 to 5cm | 30 | 28.3 ± 9.4 | 14 | 16 | 23 |  | oral | oral temperature sensor, Welch-Allyn |  |  | 0.429 | 0.359
[9 empty] | FDIR-V16 | [12 empty] | 0.237 | 0.315
[9 empty] | 70121 | [12 empty] | -0.184 | 0.283
L | L | L | L | L | L | 2021 | Stacey J. L. Sullivan [22] | hospital | NCIT1,brand unknown | forehead | 0.5 inches to 6 inches | 1113 |  | 445 | 668 | range 20.2 to 29.3 |  | oral | Welch Allyn oral thermometer (SureTemp Plus 690, Welch Allyn, San Diego, CA) |  |  | -0.23 | 0.46
[9 empty] | NCIT2,brand unknown | [12 empty] | -0.22 | 0.43
[9 empty] | NCIT3,brand unknown | [12 empty] | 0.15 | 0.42
[9 empty] | NCIT4,brand unknown | [12 empty] | -0.32 | 0.61
[9 empty] | NCIT5,brand unknown | [12 empty] | -0.87 | 0.53
[9 empty] | NCIT6,brand unknown | [12 empty] | 0.21 | 0.48
L | L | L | L | L | L | 2022 | Chakrapani Mahabala [23] | hospital | Accu DIGIT F1- BPL (India) | sublingual (oral) | 1cm | 35 | 32.8 ± 11.57 |  |  |  | 37.54±0.93 | tympanic | T-clinic TherCom cartable device (Manufacturer: Innovatec Sensing and Communication, Alicante, Spain. Model No:SN 58021315) | 0.86 |  | -0.21 | 0.48 | 38 |  | 90% | 90% | 0.96
L | L | L | L | L | L | 2020 | Srinivasan Thiagarajan [24] | hospital | Equinox Meditech Private Limited, New Delhi | forehead | 0.5 to 1cm | 250 | range 1st day of life to 12 month of age |  |  |  | 37.6±2.49 | axillary | mercury-inglass thermometer |  |  | -0.016
L | L | L | L | L | L | 2005 | Daniel K. Ng [25] | hospital | Standard ST 8812 (Standard InstrumentsCo., Hong Kong SAR, China) | forehead | 5cm | 567 | range 1 month to 18 years | 335 | 232 | 26 | range 36.0 to 41.5 | tympanic | infrared thermometer (FirstTempHGenius, Intelligent Medical Systems Inc.,Carlsbad, California) |  |  | -2.34 | 1.06 | 38 | 35.1 | 89.43% | 75.37% | 0.868 | 33.74% | 98.07%
L | L | L | L | L | L | 2016 | Hurşit Apa [26] | hospital | ThermoFlash LX-26, Visiomed SAS France, Paris/France | umbilical | 1.5cm | 100 | 56.3±50.2 months(range 1 to 168 months) | 53 | 47 | range 24 to26 | 36.8±1.03 | axillary | axillary digital thermometer (Microlife MT 3001, Microlife AG Swiss Corporation, Widnau/ Switzerland) | 0.78 |  | -0.47 | 0.65 | 38 |  | 71.1% | 95.8% | 0.93
U | H | H | L | H | H | 1992 | C. Hershler [27] | room | FirstTemp@ (Intelligent Medi- cal Systems Inc., 6339 Paseo de Lago, Carlsbad, CA 92009, USA, or Pennco, 6 South Hill Park, London NW3 2SB, UK) | skin | 1 to 3 cm | 5 | range 29 to 39 years | 5 | 0 |  | range 27.3 to 34.5 | skin | thermocouple contact sensor(Thermalert TH-5) | 0.937 |  | 0.04 | 0.05
L | L | L | L | L | L | 2020 | Gail Hayward [28] | hospital | Thermofocus | forehead |  | 401 | median 1.6 | 203 | 198 |  |  | axillary | digital axillary thermometer, Welch Allyn SureTemp |  |  | -0.14 | 0.729 | 38 |  | 29.3% | 96.9% |  | 52.2% | 92.3%
[9 empty] | Firhealth | [12 empty] | -0.16 | 0.704 |  |  | 4.9% | 98.9% |  | 33.3% | 90.1%
H | L | H | L | H | H | 2015 | Aaron J E Bach [29] | room,rest | VisioFocus 06400 Infrared Thermometer, Tecnimed Inc., Varese, Italy | skin | 60mm | 30 | 25.0±2.9 | 30 | 0 | 24.3 ± 2.5 |  | skin | two conductive devices (thermistors, iButtons)((EU-UU-VL5–0 Thermistors, Grant Instruments, Cambridge, UK; and DS1922L-F50 iButtons, Maxim Intergrated, Sunnyvale, California, USA)) |  |  | 0.34 | 0.224
[8 empty] | room,exercise | [13 empty] | -0.44 | 0.6275
H | L | H | L | H | H | 2007 | Michael J Buono [30] | environmental chamber,rest | Extech Instruments | skin | 1 cm | 6 | 25 ± 3 | 0 | 6 | 15,25,and 35 |  | skin | YSI (series 400) contact thermistors | 0.95
[8 empty] | environmental chamber,exercise | [11 empty] | 0.98
U | L | H | L | H | H | 2009 | N Ruopsa [31] | room | Raytek PhotoTemp MX6，Total Temperature Instrumentation Inc.，Williston，USA | middle finger skin | 30cm ± 3 cm | 38 |  |  |  | >31.5℃ |  | middle finger skin | wired thermometer, Datex-Engstrom, Helsinki, Finland |  |  | -0.06
[16 empty] | <31.5℃ | [5 empty] | -1.01
L | L | L | L | L | L | 2023 | Thomas Holder [32] | room | the JXB-182 Infrared Forehead Thermometer (IR gun) (Berrcom, China); the FLIR C3 Thermal Camera (IR camera) (FLIR, USA). | face | 3cm | 119 | range 20 to 59 | 50 | 69 | 20 | 36.9 ±0.37 | skin (core temperature) | SpotOn (3M, USA) |  |  | -0.19 | 0.398
[11 empty] | 6cm | [10 empty] | -0.35 | 0.403
[11 empty] | 9cm | [10 empty] | -0.51 | 0.429
L | L | L | L | L | L | 2016 | Sara Sollai [33] | hospital | Thermofocus, model 0800; Tecnimed, Varese, Italy | forehead |  | 119 healthy newborns | Mean gestational age:39 weeks + 6 days | 64 | 55 |  |  | axillary | SANITAS Hans Dislage GmbH |  |  | 0.12 | 0.444
[12 empty] | 70 preterm newborns | Mean gestational age:27 weeks + 3 days | 28 | 42 |  |  | axillary |  |  |  | 0.1 | 0.346
[12 empty] | 119 healthy newborns | Mean gestational age:39 weeks + 6 days | 64 | 55 |  |  | tympanic | Braun ThermoScan PRO 4000 |  |  | 0.02 | 0.388
[12 empty] | 70 preterm newborns | Mean gestational age:27 weeks + 3 days | 28 | 42 |  |  | tympanic |  |  |  | 0.14 | 0.276

## Reference
[1] | Teran, C.G.; Torrez-Llanos, J.; Teran-Miranda, T.E.; Balderrama, C.; Shah, N.S.; Villarroel, P. Clinical accuracy of a non-contact infrared skin thermometer in paediatric practice. Child Care Health Dev 2012, 38, 471-476, doi:10.1111/j.1365-2214.2011.01264.x.
[2] | Basak, T.; Aciksoz, S.; Tosun, B.; Akyuz, A.; Acikel, C. Comparison of three different thermometers in evaluating the body temperature of healthy young adult individuals. Int J Nurs Pract 2013, 19, 471-478, doi:10.1111/ijn.12097.
[3] | Chiappini, E.; Sollai, S.; Longhi, R.; Morandini, L.; Laghi, A.; Osio, C.E.; Persiani, M.; Lonati, S.; Picchi, R.; Bonsignori, F.; et al. Performance of non-contact infrared thermometer for detecting febrile children in hospital and ambulatory settings. J Clin Nurs 2011, 20, 1311-1318, doi:10.1111/j.1365-2702.2010.03565.x.
[4] | Williams, E.M.; Heusch, A.I.; McCarthy, P.W. Thermal screening of facial skin arterial hot spots using non-contact infrared radiometry. Physiol Meas 2008, 29, 341-348, doi:10.1088/0967-3334/29/3/005.
[5] | Costanzo, S.; Flores, A. A Non-Contact Integrated Body-Ambient Temperature Sensors Platform to Contrast COVID-19. Electronics 2020, 9, doi:10.3390/electronics9101658.
[6] | Kameda, N. Clinical accuracy of non-contact forehead infrared thermometer and infrared tympanic thermometer in postoperative adult patients: A comparative study. J Perioper Pract 2022, 32, 142-148, doi:10.1177/17504589211022314.
[7] | Erdem, N.; Demirdag, T.B.; Tezer, H.; Cura-Yayla, B.C.; Baran-Aksakal, F.N.; Tapisiz, A.; Derinoz, O.; Okur, A.; Pinarli, F.G.; Kocak, U.; et al. The comparison and diagnostic accuracy of different types of thermometers. Turk J Pediatr 2021, 63, 434-442, doi:10.24953/turkjped.2021.03.010.
[8] | Hakan, N.; Okumuş, N.; Aydın, M.; Küçüközkan, T.; Tuygun, N.; Zenciroğlu, A. The comparison of temporal temperature measurement method by non-contact infrared thermometer with other body temperature measurement methods. Journal of Dr Behcet Uz Children s Hospital 2017, doi:10.5222/buchd.2017.141.
[9] | Franconi, I.; La Cerra, C.; Marucci, A.R.; Petrucci, C.; Lancia, L. Digital Axillary and Non-Contact Infrared Thermometers for Children. Clin Nurs Res 2018, 27, 180-190, doi:10.1177/1054773816676538.
[10] | Huang, F.; Magnin, C.; Brouqui, P. Ingestible sensors correlate closely with peripheral temperature measurements in febrile patients. J Infect 2020, 80, 161-166, doi:10.1016/j.jinf.2019.11.003.
[11] | Lai, F.; Li, X.; Liu, T.; Wang, X.; Wang, Q.; Chen, S.; Wei, S.; Xiong, Y.; Hou, Q.; Zeng, X.; et al. Optimal diagnostic fever thresholds using non-contact infrared thermometers under COVID-19. Front Public Health 2022, 10, 985553, doi:10.3389/fpubh.2022.985553.
[12] | Pana, B.C.; Lopes, H.; Furtunescu, F.; Franco, D.; Rapcea, A.; Stanca, M.; Tanase, A.; Colita, A. Real-World Evidence: The Low Validity of Temperature Screening for COVID-19 Triage. Front Public Health 2021, 9, 672698, doi:10.3389/fpubh.2021.672698.
[13] | Lai, F.; Li, X.; Wang, Q.; Luo, Y.; Wang, X.; Huang, X.; Zhang, J.; Peng, J.; Wang, Q.; Fan, L.; et al. Reliability of Non-Contact Infrared Thermometers for Fever Screening Under COVID-19. Risk Manag Healthc Policy 2022, 15, 447-456, doi:10.2147/RMHP.S357567.
[14] | Kocoglu, H.; Goksu, S.; Isik, M.; Akturk, Z.; Bayazit, Y.A. Infrared tympanic thermometer can accurately measure the body temperature in children in an emergency room setting. Int J Pediatr Otorhinolaryngol 2002, 65, 39-43, doi:10.1016/s0165-5876(02)00129-5.
[15] | Moran-Navarro, R.; Courel-Ibanez, J.; Martinez-Cava, A.; Conesa-Ros, E.; Sanchez-Pay, A.; Mora-Rodriguez, R.; Pallares, J.G. Validity of Skin, Oral and Tympanic Temperatures During Exercise in the Heat: Effects of Wind and Sweat. Ann Biomed Eng 2019, 47, 317-331, doi:10.1007/s10439-018-02115-x.
[16] | Kim, Y.M.; Jang, M.R.; Moon, J.R.; Park, G.; An, Y.J.; Seo, J.M. Clinical Accuracy of Non-Contact Forehead Infrared Thermometer Measurement in Children: An Observational Study. Children (Basel) 2022, 9, doi:10.3390/children9091389.
[17] | Ray, P.P. An IR Sensor Based Smart System to Approximate Core Body Temperature. J Med Syst 2017, 41, 123, doi:10.1007/s10916-017-0770-z.
[18] | Ajcevic, M.; Buoite Stella, A.; Furlanis, G.; Caruso, P.; Naccarato, M.; Accardo, A.; Manganotti, P. A Novel Non-Invasive Thermometer for Continuous Core Body Temperature: Comparison with Tympanic Temperature in an Acute Stroke Clinical Setting. Sensors (Basel) 2022, 22, doi:10.3390/s22134760.
[19] | Goh, N.W.; Poh, J.J.; Yeo, J.Y.; Aw, B.J.; Lai, S.C.; Cheng, J.J.W.; Tan, C.Y.L.; Gan, S.K. Design and Development of a Low Cost, Non-Contact Infrared Thermometer with Range Compensation. Sensors (Basel) 2021, 21, doi:10.3390/s21113817.
[20] | Rinanto, N.; Kuo, C.-H. PCA-ANN Contactless Multimodality Sensors for Body Temperature Estimation. IEEE Transactions on Instrumentation and Measurement 2021, 70, 1-16, doi:10.1109/tim.2021.3112778.
[21] | Mah, A.J.; Ghazi Zadeh, L.; Khoshnam Tehrani, M.; Askari, S.; Gandjbakhche, A.H.; Shadgan, B. Studying the Accuracy and Function of Different Thermometry Techniques for Measuring Body Temperature. Biology (Basel) 2021, 10, doi:10.3390/biology10121327.
[22] | Sullivan, S.J.L.; Rinaldi, J.E.; Hariharan, P.; Casamento, J.P.; Baek, S.; Seay, N.; Vesnovsky, O.; Topoleski, L.D.T. Clinical evaluation of non-contact infrared thermometers. Sci Rep 2021, 11, 22079, doi:10.1038/s41598-021-99300-1.
[23] | Mahabala, C.; Dakappa, P.H.; Gupta, A.R. A novel method for measuring sublingual temperature using conventional non-contact forehead thermometer. F1000Res 2022, 11, 13, doi:10.12688/f1000research.74876.3.
[24] | Thiagarajan, S.; Balaji, R.; Pothapregada, S. Non-Contact Infrared Thermometry in Febrile Infants. Indian Pediatrics. 2020, 57, 857–858.
[25] | Ng, D.K.; Chan, C.H.; Lee, R.S.; Leung, L.C. Non-contact infrared thermometry temperature measurement for screening fever in children. Ann Trop Paediatr 2005, 25, 267-275, doi:10.1179/146532805X72412.
[26] | Apa, H.; Gozmen, S.; Keskin-Gozmen, S.; Aslan, F.; Bayram, N.; Devrim, I. Clinical accuracy of non-contact infrared thermometer from umbilical region in children: A new side. Turk J Pediatr 2016, 58, 180-186, doi:10.24953/turkjped.2016.02.009.
[27] | Hershler, C.; Conine, T.A.; Nunn, A.; Hannay, M. Assessment of an infra-red non-contact sensor for routine skin temperature monitoring: a preliminary study. J Med Eng Technol 1992, 16, 117-122, doi:10.3109/03091909209021973.
[28] | Hayward, G.; Verbakel, J.Y.; Ismail, F.A.; Edwards, G.; Wang, K.; Fleming, S.; Holtman, G.A.; Glogowska, M.; Morris, E.; Curtis, K.; et al. Non-contact infrared versus axillary and tympanic thermometers in children attending primary care: a mixed-methods study of accuracy and acceptability. Br J Gen Pract 2020, 70, e236-e244, doi:10.3399/bjgp20X708845.
[29] | Bach, A.J.; Stewart, I.B.; Disher, A.E.; Costello, J.T. A comparison between conductive and infrared devices for measuring mean skin temperature at rest, during exercise in the heat, and recovery. PLoS One 2015, 10, e0117907, doi:10.1371/journal.pone.0117907.
[30] | Buono, M.J.; Jechort, A.; Marques, R.; Smith, C.; Welch, J. Comparison of infrared versus contact thermometry for measuring skin temperature during exercise in the heat. Physiol Meas 2007, 28, 855-859, doi:10.1088/0967-3334/28/8/008.
[31] | Ruopsa, N.; Kujala, S.; Kaarela, O.; Ohtonen, P.; Ryhanen, J. Wireless infrared thermometer in the follow-up of finger temperatures. J Hand Surg Eur Vol 2009, 34, 526-529, doi:10.1177/1753193409102456.
[32] | Holder, T.; Hooper, F.S.W.; Yates, D.; Tse, Z.; Patil, S.; Moussa, A.; Batten, L.; Radhakrishnan, V.; Allison, M.; Hewitt, C.; et al. Clinical accuracy of infrared temperature measurement devices: a comparison against non-invasive core-body temperature. Clin Med (Lond) 2023, 23, 157-163, doi:10.7861/clinmed.2022-0252.
[33] | Sollai, S.; Dani, C.; Berti, E.; Fancelli, C.; Galli, L.; de Martino, M.; Chiappini, E. Performance of a non-contact infrared thermometer in healthy newborns. BMJ Open 2016, 6, e008695, doi:10.1136/bmjopen-2015-008695.