## Paper Info
Vancomycin & Tobramycin
Flow Rate | 0.3mL/min |  |  |  | Lowest | Highest
Column Oven | 40°C |  |  | Detection Limits | 0.98 | 250 | µg/mL
Gradient | 2.00 min | 70% A | 30% B
| 5.00 min | 5% A | 95% B
| 6.00 min | 5% A | 95% B |  | Below limits of detection
| 6.10 min | 90% A | 5% B |  | No sample
Stop time | 10.00 min
Post time | 5.00 min
CSBv+tSpaceru data was not used because it was clinically irrelevant
Method Development | V&T
Injection Source | Als
Injection Location | Als
Binary Pump | G7112B
Flow Rate | 0.300 mL/min
High Pressure Limit | 400 bar
Stoptime | 10.00 min
Posttime | 5.00 min
Mobile Phase A | H2O + 0.1% HFBA
Mobile Phase B | ACN + 0.1% HFBA
Sampler | G7129A
Draw Speed | 200 µL/min
Eject Speed | 400 µL/min
Wait time after draw | 1.2 secs
Needle Wash | Yes
Column Oven | G7130A
Temperature set | 40°C
Not Ready Limit Value | 0.8°C
Mass Spectrometer Detector
Ionization mode | API-ES
Polarity | Positive
SIM Parameters
SIM Ion | 468.2 T & 724.8 V
Fragmentor | 130
Gain EMV | 1
Actual dwell | 294
Spray Chamber
Gas Temp | 345°C
Drying Gas | 9.0 L/min
Neb Pressure | 50 psig
Vcap (Positive) | 3500V
Vcap (Negative) | 3000V
General Calibration Settings
Curve type | Quadratic
Origin | Ignored
Weight | Equal
Machine
Agilent Infinity II 1260 Lab LC/MSD
Column
Poroshell 120
SB-C18 2.7µm, 2.1 x 100mm
PN: | 685775-902(T)
SN: | USCFH07050
LN: | B18178

## CSBv+tSpaceru
This data was not used because it was clinically irrelevant
| TOB Concentration (µg/mL) |  |  |  |  | VANC Concentration (µg/mL) | [6 empty] | Spacer (no antibiotic) + Beads (1g Vancomycin & 240mg Tobramycin)
| G1 | G1Br | G1C | G1E | G1F |  | G1 | G1Br | G1C | G1E | G1F |  | Averages over all repeats |  |  | Std Dev over all repeats
Blank | 2.241 |  |  |  |  | Blank | 0
0 | 602.083 |  |  |  |  | 0 | 364.565 | [5 empty] | Timepoint | Tob | Vanc |  | Timepoint | Tob | Vanc
0 | 603.309 |  |  |  |  | 0 | 356.007 | [5 empty] | 0 | 601.2443 | 358.4497 |  | 0 | 2.588 | 5.3316
0 | 598.341 |  |  |  |  | 0 | 354.777 | [5 empty] | 0.05 | 472.9425 | 610.9778 |  | 0.05 | 656.1543 | 679.6126
Blank | 0 | 0.693 |  |  |  | Blank | 0 | 0 |  |  |  |  | 0.1 | 418.8547 | 770.4378 |  | 0.1 | 475.4607 | 744.0746
5 |  | 53.276 | 124.124 | 171.347 | 1554.985 | 5 |  | 0 | 209.329 | 438.982 | 1756.628 |  | 0.15 | 316.0311 | 712.2167 |  | 0.15 | 331.9677 | 793.7121
5 |  | 53.642 | 124.014 | 150.533 | 1567.813 | 5 |  | 0 | 206.651 | 628.151 | 1739.078 |  | 0.2 | 221.7635 | 711.4407 |  | 0.2 | 250.0197 | 706.9879
5 |  | 53.641 | 124.455 | 142.909 | 1554.571 | 5 |  | 0 | 197.532 | 622.532 | 1532.851 |  | 0.25 | 155.0757 | 501.0972 |  | 0.25 | 178.753 | 502.0618
blank |  | 0.716 | 0 | 0 | 0 | BLANK |  | 0 | 0 | 0 | 0 |  | 0.3 | 103.3566 | 396.9729 |  | 0.3 | 139.9629 | 340.2595
10 |  | 17.778 | 118.83 | 355.403 | 1161.33 | 10 |  | 160.774 | 190.851 | 1080.377 | 1884.241 |  | 0.35 | 78.2042 | 299.0677 |  | 0.35 | 103.6376 | 208.1949
10 |  | 17.598 | 119.529 | 355.784 | 1167.467 | 10 |  | 109.26 | 181.412 | 753.383 | 1866.969 |  | 0.4 | 68.8135 | 242.0782 |  | 0.4 | 101.3942 | 173.7031
10 |  | 18.017 | 121.162 | 370.965 | 1202.393 | 10 |  | 157.536 | 167.375 | 773.829 | 1919.247 |  | 0.45 | 48.7231 | 137.6192 |  | 0.45 | 65.6674 | 67.1113
blank |  | 0.677 | 0 | 0 | 0 | BLANK |  | 0 | 0 | 0 | 0 |  | 0.5 | 41.6942 | 106.8276 |  | 0.5 | 53.3187 | 29.2879
15 | 190.918 | 14.896 | 26.641 | 480.451 | 857.017 | 15 | 140.981 | 133.593 | 220.484 | 1173.443 | 2064.865 |  | 0.55 | 50.294 | 122.7218 |  | 0.55 | 45.6403 | 49.1496
15 | 187.306 | 14.143 | 26.768 | 508.431 | 854.026 | 15 | 140.383 | 93.481 | 221.831 | 892.975 | 2100.948 |  | 1 | 175.4629 | 150.5792 |  | 1 | 217.6312 | 121.0169
15 | 188.438 | 13.938 | 26.765 | 494.034 | 856.695 | 15 | 141.339 | 134.887 | 221.702 | 861.583 | 2140.755 |  | 2 | 91.9715 | 151.4665 |  | 2 | 136.4302 | 73.9977
Blank | 0 | 0.66 | 0.67 | 0 | 0 | Blank | 0 | 0 | 0 | 0 | 0 |  | 3 | 106.9541 | 203.5295 |  | 3 | 158.4374 | 116.2477
20 |  | 14.313 | 19.07 | 257.057 | 589.118 | 20 |  | 127.958 | 155.823 | 957.601 | 1795.711 |  | 4 | 126.0156 | 234.3955 |  | 4 | 149.4501 | 103.352
20 |  | 14.19 | 18.937 | 251.926 | 609.306 | 20 |  | 129.806 | 154.768 | 678.535 | 1774.728 |  | 5 | 146.9142 | 328.0634 |  | 5 | 154.1059 | 250.6873
20 |  | 14.269 | 18.647 | 249.289 | 605.04 | 20 |  | 128.496 | 155.144 | 679.797 | 1798.921 |  | 6 | 163.1586 | 369.5431 |  | 6 | 153.4683 | 246.8369
blank |  | 0.664 | 0 | 0 | 0 | Blank |  | 0 | 0 | 0 | 0 |  | 7 | 135.7531 | 219.632 |  | 7 | 142.7571 | 174.2662
25 |  | 12.54 | 15.089 | 161.325 | 418.879 | 25 |  | 83.831 | 95.414 | 560.329 | 1224.865 |  | 8 | 272.8827 | 427.1747 |  | 8 | 272.7417 | 208.1215
25 |  | 12.662 | 14.748 | 153.992 | 431.211 | 25 |  | 83.243 | 93.013 | 537.013 | 1304.677 |  | 12 | 199.901 | 787.8426 |  | 12 | 226.3845 | 1181.2723
25 |  | 12.515 | 14.92 | 165.312 | 447.715 | 25 |  | 83.735 | 124.844 | 546.027 | 1276.176 |  | 18 | 245.2257 | 905.776 |  | 18 | 230.8607 | 1334.8249
blank |  | 0.647 | 0.645 | 0 | 0 | BLANK |  | 0 | 0 | 0 | 0 |  | 24 | 221.1083 | 820.9921 |  | 24 | 206.749 | 1192.5337
30 |  | 12.71 | 15.174 | 77.462 | 341.836 | 30 |  | 114.464 | 94.858 | 371.443 | 904.566 |  | 48 | 206.4304 | 1009.2498 |  | 48 | 250.3252 | 1301.9443
30 |  | 12.796 | 15.293 | 46.137 | 340.097 | 30 |  | 86.281 | 129.755 | 519.057 | 902.319 | D5 | 120 | 170.2972 | 895.919 | D5 | 120 | 265.3819 | 1321.6878
30 |  | 12.644 | 17.183 | 31.464 | 317.483 | 30 |  | 114.688 | 127.736 | 510.814 | 887.694 | D8 | 192 | 137.8056 | 743.6812 | D8 | 192 | 246.0014 | 1153.0864
blank |  | 0 | 0.644 | 0 | 0 | Blank |  | 0 | 0 | 0 | 0 | D11 | 264 | 215.9889 | 1121.533 | D11 | 264 | 321.8494 | 1678.9868
35 |  | 13.459 | 16.917 | 0 | 244.985 | 35 |  | 120.456 | 130.599 | 326.431 | 567.502 | D14 | 336 | 94.8634 | 656.3902 | D14 | 336 | 140.1469 | 981.0245
35 |  | 13.654 | 16.832 | 70.012 | 250.054 | 35 |  | 88.094 | 127.514 | 406.216 | 645.621 | D17 | 408 | 58.36 | 547.4539 | D17 | 408 | 85.4735 | 817.4663
35 |  | 13.553 | 17.154 | 33.762 | 248.068 | 35 |  | 119.649 | 124.648 | 366.51 | 565.572 | D20 | 480 | 53.8224 | 379.7247 | D20 | 480 | 78.5999 | 566.4492
blank |  | 0.669 | 0 | 0 | 0 | Blank |  | 0 | 0 | 0 | 0 | D23 | 552 | 58.716 | 240.4486 | D23 | 552 | 85.9861 | 358.4846
40 |  | 13.654 | 16.656 | 8.251 | 232.628 | 40 |  | 116.692 | 93.92 | 239.98 | 508.954 | D26 | 624 | 35.8943 | 167.6616 | D26 | 624 | 51.757 | 249.6295
40 |  | 13.623 | 16.811 | 8.112 | 239.338 | 40 |  | 116.216 | 123.792 | 236.272 | 518.323 | D29 | 696 | 7.7324 | 82.2697 | D29 | 696 | 9.4055 | 122.9012
40 |  | 13.48 | 16.585 | 8.031 | 238.593 | 40 |  | 87.448 | 122.412 | 219.424 | 521.505 | D32 | 768 | 6.829 | 54.56 | D32 | 768 | 8.069 | 81.3662
blank |  | 0.641 | 0.642 | 0 | 0 | Blank |  | 0 | 0 | 0 | 0 | D35 | 840 | 6.4184 | 33.3069 | D35 | 840 | 7.4632 | 49.2426
45 |  | 13.777 | 15.443 | 8.123 | 156.556 | 45 |  | 115.579 | 111.397 | 130.652 | 251.693 | D38 | 912 | 4.54 | 15.0457 | D38 | 912 | 3.4062 | 20.4637
45 |  | 13.969 | 15.539 | 8.075 | 157.751 | 45 |  | 115.701 | 83.612 | 36.846 | 233.439 | D41 | 984 | 3.4053 | 8.4943 | D41 | 984 | 2.2014 | 11.5368
45 |  | 13.753 | 15.39 | 8.063 | 158.238 | 45 |  | 116.919 | 111.59 | 104.531 | 239.471
blank |  | 0 | 0.648 | 0 | 0 | Blank |  | 0 | 0 | 0 | 0
50 |  | 13.39 | 15.446 | 7.96 | 128.955 | 50 |  | 114.157 | 83.473 | 83.516 | 145.349
50 |  | 13.455 | 15.446 | 7.933 | 130.09 | 50 |  | 82.947 | 108.736 | 81.674 | 153.473
50 |  | 13.447 | 15.427 | 7.88 | 130.901 | 50 |  | 84.655 | 109.582 | 80.354 | 154.015
blank |  | 0.648 | 0.642 | 0 | 0 | Blank |  | 0 | 0 | 0 | 0
55 |  | 12.966 | 15.951 | 50.451 | 119.741 | 55 |  | 83.898 | 116.333 | 203.556 | 89.96
55 |  | 13.582 | 16.327 | 50.704 | 121.983 | 55 |  | 83.765 | 115.299 | 196.083 | 91.459
55 |  | 13.486 | 15.333 | 49.913 | 123.091 | 55 |  | 83.045 | 114.766 | 204.887 | 89.611
blank |  | 0 | 0 | 0 | 0 | Blank |  | 0 | 0 | 0 | 0
60 | 591.288 | 16.204 | 110.953 | 55.454 | 104.089 | 60 | 343.171 | 132.431 | 61.66 | 214.372 | 14.448
60 | 591.8 | 16.119 | 110.266 | 56.284 | 103.572 | 60 | 339.186 | 133.953 | 59.197 | 208.573 | 7.441
60 | 587.874 | 16.192 | 110.387 | 56.524 | 104.938 | 60 | 341.517 | 131.725 | 57.18 | 206.169 | 7.665
Blank | 0 | 0.656 | 0 | 0 | 0 | Blank | 0 | 0 | 0 | 0 | 0
2 | 350.859 | 15.415 | 19.184 | 60.365 | 11.343 | 2 | 251.416 | 128.891 | 134.212 | 202.578 | 59.154
2 | 354.418 | 15.748 | 19.37 | 60.288 | 11.589 | 2 | 261.883 | 93.18 | 97.878 | 200.621 | 58.15
2 | 354.459 | 15.646 | 19.496 | 59.851 | 11.541 | 2 | 268.633 | 127.481 | 133.178 | 195.724 | 59.019
Blank | 0 | 0 | 0 | 0 | 0 | Blank | 0 | 0 | 0 | 0 | 0
3 | 409.509 | 27.117 | 27.793 | 51.311 | 15.761 | 3 | 372.176 | 291.57 | 207.39 | 163.93 | 84.293
3 | 411.87 | 27.345 | 27.93 | 52.806 | 15.602 | 3 | 369.351 | 0 | 210.058 | 162.812 | 82.756
3 | 415.087 | 26.786 | 27.648 | 52.101 | 15.646 | 3 | 380.578 | 275.104 | 203.107 | 164.636 | 85.181
Blank | 0 | 0.707 | 0 | 0 | 0 | Blank | 0 | 0 | 0 | 0 | 0
4 | 403.145 | 129.16 | 29.908 | 48.168 | 17.888 | 4 | 394.729 | 280.995 | 222.555 | 177.797 | 100.526
4 | 404.258 | 129.646 | 29.666 | 49.107 | 17.652 | 4 | 411.9 | 267.834 | 225.229 | 185.613 | 99.516
4 | 404.707 | 129.924 | 29.857 | 49.402 | 17.746 | 4 | 393.348 | 256.778 | 223.048 | 174.435 | 101.629
Blank | 0 | 0 | 0 | 0 | 0 | Blank | 0 | 0 | 0 | 0 | 0
5 | 415.647 | 0 | 155.28 | 27.717 | 18.62 | 5 | 450.319 | 0 | 529.568 | 87.828 | 104.333
5 | 415.051 | 173.022 | 161.584 | 27.452 | 18.468 | 5 | 445.978 | 704.879 | 523.563 | 91.413 | 104.019
5 | 416.814 | 171.8 | 156.547 | 27.287 | 18.424 | 5 | 435.076 | 716.06 | 530.866 | 92.977 | 104.072
Blank | 0 | 0 | 0 | 0 | 0 | Blank | 0 | 0 | 0 | 0 | 0
6 | 427.106 | 171.689 | 158.709 | 41.683 | 11.74 | 6 | 480.402 | 691.299 | 503.672 | 112.785 | 72.145
6 | 431.28 | 172.26 | 161.06 | 40.701 | 11.618 | 6 | 483.57 | 682.049 | 492.662 | 104.959 | 71.72
6 | 435.93 | 170.706 | 160.263 | 41.03 | 11.604 | 6 | 491.354 | 681.528 | 489.136 | 112.887 | 72.979
Blank | 0 | 0 | 0 | 0 | 0 | Blank | 0 | 0 | 0 | 0 | 0
7 | 397.008 | 50.236 | 135.521 | 78.943 | 13.246 | 7 | 436.171 | 0 | 375.369 | 195.696 | 83.209
7 | 402.604 | 50.352 | 135.96 | 79.787 | 13.27 | 7 | 432.341 | 0 | 390.157 | 193.162 | 82.864
7 | 398.945 | 49.451 | 138.083 | 79.603 | 13.287 | 7 | 454.623 | 0 | 367.721 | 200.98 | 82.187
Blank | 0 | 0.716 | 0 | 0 | 0 | Blank | 0 | 0 | 0 | 0 | 0
8 | 721.263 | 139.733 | 127.59 | 97.176 |  | 8 | 763.225 | 422.022 | 275.197 | 235.248
8 | 733.78 | 142.113 | 131.109 | 96.321 |  | 8 | 749.475 | 435.407 | 288.085 | 248.678
8 | 717.577 | 142.868 | 130.265 | 94.797 |  | 8 | 743.744 | 420.161 | 283.82 | 261.034
Blank | 0 | 0 | 0 | 0 |  | Blank | 0 | 0 | 0 | 0
12 | 629.427 | 126.843 | 117.358 | 108.479 | 16.545 | 12 | 3048.449 | 321.192 | 179.185 | 350.301 | 147.252
12 | 635.706 | 128.016 | 116.515 | 107.388 | 16.477 | 12 | 3057.594 | 328.954 | 173.632 | 0 | 146.668
12 | 625 | 128.293 | 117.093 | 108.946 | 16.429 | 12 | 3081.227 | 326.318 | 172.214 | 334.783 | 149.87
Blank | 0 | 0 | 0 | 0 | 0 | Blank | 0 | 0 | 0 | 0 | 0
18 | 640.698 | 111.695 | 114.063 | 130.355 |  | 18 | 3120.737 | 296.948 | 221.289 | 0
18 | 620.365 | 113.032 | 109.846 | 128.843 |  | 18 | 3098.727 | 299.671 | 207.275 | 0
18 | 622.298 | 113.581 | 109.973 | 127.96 |  | 18 | 3114.271 | 302.828 | 207.566 | 0
Blank | 0 | 0 | 0 | 0 |  | Blank | 0 | 0 | 0 | 0
24 | 615.129 | 109.41 | 106.38 | 163.969 | 106.299 | 24 | 3075.387 | 336.171 | 228.276 | 0 | 436.218
24 | 622.821 | 112.414 | 107.507 | 162.259 | 106.687 | 24 | 3108.419 | 343.451 | 234.242 | 0 | 427.295
24 | 617.124 | 112.093 | 106.626 | 160.746 | 107.161 | 24 | 3137.658 | 333.17 | 229.579 | 0 | 425.016
Blank | 0 | 0 | 0 | 0 | 0 | Blank | 0 | 0 | 0 | 0 | 0
48 | 611.255 | 3.629 | 129.085 |  | 78.126 | 48 | 3190.946 | 213.28 | 201.814 |  | 460.835
48 | 622.258 | 3.607 | 130.298 |  | 78.553 | 48 | 3169.247 | 201.308 | 203.296 |  | 461.878
48 | 609.042 | 3.515 | 129.775 |  | 78.022 | 48 | 3121.313 | 219.95 | 201.031 |  | 466.099
Blank | 0 | 0 | 0 |  | 0 | Blank | 0 | 0 | 0 |  | 0
D5 | 601.868 | 1.371 | 1.309 |  | 70.848 | D5 | 3034.34 | 53.524 | 43.853 |  | 409.34
D5 | 606.139 | 1.347 | 1.318 |  | 70.731 | D5 | 3102.533 | 53.485 | 51.618 |  | 411.027
D5 | 614.786 | 1.337 | 1.29 |  | 71.223 | D5 | 3080.939 | 45.009 | 52.044 |  | 413.316
Blank | 0 | 0 | 0 |  | 0 | Blank | 0 | 0 | 0 |  | 0
D8 | 542.758 | 1.38 | 1.323 |  | 2.849 | D8 | 2623.533 | 2.213 | 1.84 |  | 329.293
D8 | 547.621 | 1.376 | 1.326 |  | 2.772 | D8 | 2684.559 | 2.17 | 1.835 |  | 327.816
D8 | 546.863 | 1.372 | 1.314 |  | 2.713 | D8 | 2616.826 | 2.197 | 1.817 |  | 330.075
Blank | 0 | 0 | 0 |  | 0 | Blank | 0 | 0 | 0 |  | 0
D11 | 639.112 | 1.438 | 1.45 |  |  | D11 | 3317.609 | 2.693 | 1.983
D11 | 644.213 | 1.441 | 1.422 |  |  | D11 | 3366.003 | 2.666 | 1.839
D11 | 651.974 | 1.447 | 1.403 |  |  | D11 | 3396.463 | 2.657 | 1.884
blank | 0 | 0 | 0 |  |  | blank | 0 | 0 | 0
D14 | 286.815 | 1.438 | 1.483 |  |  | D14 | 1947.948 | 2.759 | 2.149
D14 | 282.813 | 1.442 | 1.442 |  |  | D14 | 1994.112 | 2.755 | 2.11
D14 | 275.431 | 1.442 | 1.465 |  |  | D14 | 1950.868 | 2.706 | 2.105
Blank | 2.146 | 0 | 0 |  |  | Blank | 0 | 0 | 0
D17 | 174.827 | 1.384 | 1.367 |  |  | D17 | 1645.553 | 2.869 | 2.116
D17 | 169.599 | 1.413 | 1.375 |  |  | D17 | 1640.745 | 2.857 | 2.08
D17 | 172.508 | 1.42 | 1.347 |  |  | D17 | 1625.864 | 2.901 | 2.1
Blank | 2.1 | 0 | 0 |  |  | Blank | 0 | 0 | 0
D20 | 157.942 | 1.483 | 1.378 |  |  | D20 | 1122.584 | 2.424 | 1.788
D20 | 161.594 | 1.474 | 1.415 |  |  | D20 | 1143.098 | 2.442 | 1.767
D20 | 156.284 | 1.458 | 1.374 |  |  | D20 | 1139.184 | 2.443 | 1.792
Blank | 2.093 | 0 | 0 |  |  | Blank | 0 | 0 | 0
D23 | 176.283 | 1.38 | 1.403 |  |  | D23 | 712.268 | 1.842 | 1.137
D23 | 173.535 | 1.381 | 1.415 |  |  | D23 | 724.783 | 1.822 | 1.15
D23 | 170.221 | 1.404 | 1.422 |  |  | D23 | 718.178 | 1.763 | 1.094
Blank | 2.129 | 0 | 0 |  |  | Blank | 0 | 0 | 0
D26 | 105.911 | 1.378 | 1.406 |  |  | D26 | 487.699 | 1.556 | 1.112
D26 | 105.615 | 1.389 | 1.378 |  |  | D26 | 512.865 | 1.521 | 1.025
D26 | 103.163 | 1.398 | 1.411 |  |  | D26 | 500.621 | 1.491 | 1.064
Blank | 2.108 | 0 | 0 |  |  | Blank | 0 | 0 | 0
D29 | 20.59 | 1.402 | 1.466 |  |  | D29 | 252.715 | 0.544 | 0.146
D29 | 20.313 | 1.471 | 1.48 |  |  | D29 | 242.48 | 0.559 | 0.164
D29 | 19.91 | 1.476 | 1.484 |  |  | D29 | 243.084 | 0.578 | 0.157
Blank | 1.972 | 0 | 0 |  |  | Blank | 0 | 0 | 0
D32 | 17.836 | 1.449 | 1.42 |  |  | D32 | 163.687 | 0.554 | 0.056
D32 | 17.238 | 1.434 | 1.445 |  |  | D32 | 163.281 | 0.545 | 0.074
D32 | 17.683 | 1.491 | 1.465 |  |  | D32 | 162.172 | 0.591 | 0.08
Blank | 0 | 0 | 0 |  |  | Blank | 0 | 0 | 0
D35 | 16.794 | 1.483 | 1.445 |  |  | D35 | 99.782 | 0.72 | 0.239
D35 | 16.049 | 1.468 | 1.427 |  |  | D35 | 99.139 | 0.683 | 0.263
D35 | 16.255 | 1.417 | 1.428 |  |  | D35 | 97.96 | 0.718 | 0.258
Blank | 0 | 0 | 0 |  |  | Blank | 0 | 0 | 0
D38 | 9.029 | 1.596 | 3.08 |  |  | D38 | 42.494 | 0.988 | 1.81
D38 | 9.103 | 1.549 | 3.015 |  |  | D38 | 42.046 | 0.958 | 1.859
D38 | 8.878 | 1.609 | 3.001 |  |  | D38 | 42.437 | 0.99 | 1.829
BLANK | 0 | 0 | 0 |  |  | BLANK | 0 | 0 | 0
D41 | 6.269 | 1.286 | 2.703 |  |  | D41 | 23.871 | 0.458 | 1.211
D41 | 6.152 | 1.289 | 2.705 |  |  | D41 | 23.791 | 0.439 | 1.186
D41 | 6.249 | 1.283 | 2.712 |  |  | D41 | 23.949 | 0.381 | 1.163
Blank | 0 | 0 | 0 |  |  | Blank | 0 | 0 | 0

## CSBv+tplusSpacerv+t
| TOB Concentration (µg/mL) | [5 empty] | VANC Concentration (µg/mL) |  |  |  |  | Spacer (2g Vancomycin & 2g Tobramycin) + Beads (1g Vancomycin & 240mg Tobramycin)
| G2 | G2B | G2C | G2D | G2E |  | G2 | G2B | G2C | G2D | G2E |  | Averages over all repeats |  |  |  | Std Dev over all repeats
Blank | 0 |  |  |  |  | Blank | 0
0 | 1010.898 |  |  |  |  | 0 | 542.214 | [5 empty] | Timepoint | Tob | Vanc |  | Timepoint | Tob | Vanc
0 | 1030.765 |  |  |  |  | 0 | 543.327 | [5 empty] | 0 | 1020.852 | 544.9723 |  | 0 | 9.9336 | 3.8541
0 | 1020.893 |  |  |  |  | 0 | 549.376 | [5 empty] | 0.05 | 619.6404 | 604.0312 |  | 0.05 | 764.267 | 580.6081
Blank | 0 |  |  |  |  | Blank | 0 | [5 empty] | 0.1 | 437.238 | 605.7864 |  | 0.1 | 530.4905 | 637.2313
0.05 |  | 87.611 | 67.655 | 462.554 | 1851.237 | 0.05 |  | 115.643 | 155.372 | 725.404 | 1480.395 |  | 0.15 | 308.7045 | 579.1191 |  | 0.15 | 334.3187 | 648.9855
0.05 |  | 84.617 | 67.146 | 460.77 | 1881.409 | 0.05 |  | 114.653 | 99.547 | 727.889 | 1489.326 |  | 0.2 | 259.4833 | 595.6694 |  | 0.2 | 282.2049 | 703.9008
0.05 |  | 84.669 | 67.245 | 485.498 | 1835.274 | 0.05 |  | 114.283 | 100.024 | 690.439 | 1435.4 |  | 0.25 | 195.5386 | 442.6089 |  | 0.25 | 199.5831 | 495.9144
Blank |  | 1.911 | 1.732 | 0 | 0 | Blank |  | 0 | 0 | 0 | 0 |  | 0.3 | 154.2626 | 332.9381 |  | 0.3 | 156.3352 | 361.3886
0.1 |  | 38.578 | 26.573 | 398.038 | 1268.835 | 0.1 |  | 104.803 | 83.678 | 689.553 | 1545.515 |  | 0.35 | 117.128 | 225.6928 |  | 0.35 | 114.1071 | 239.1442
0.1 |  | 38.079 | 27.054 | 401.829 | 1272.621 | 0.1 |  | 105.127 | 84.057 | 690.33 | 1575.516 |  | 0.4 | 93.8513 | 144.5689 |  | 0.4 | 90.2023 | 135.6418
0.1 |  | 38.42 | 26.831 | 423.884 | 1286.114 | 0.1 |  | 103.913 | 0.468 | 703.562 | 1582.915 |  | 0.45 | 71.1347 | 88.961 |  | 0.45 | 64.9519 | 83.6009
Blank |  | 1.805 | 1.694 | 0 | 0 | Blank |  | 0 | 0 | 0 | 0 |  | 0.5 | 60.3213 | 54.9152 |  | 0.5 | 56.8629 | 51.2694
0.15 | 227.551 | 29.192 | 19.41 | 371.459 | 890.426 | 0.15 | 327.014 | 94.557 | 73.845 | 667.63 | 1731.224 |  | 0.55 | 43.256 | 55.1833 |  | 0.55 | 47.2814 | 42.2078
0.15 | 227.123 | 29.819 | 19.188 | 368.852 | 894.671 | 0.15 | 314.704 | 89.615 | 72.872 | 642.651 | 1768.111 |  | 1 | 228.8398 | 125.4572 |  | 1 | 397.5795 | 207.4369
0.15 | 223.422 | 29.286 | 19.13 | 367.212 | 913.826 | 0.15 | 332.871 | 89.849 | 71.861 | 620.444 | 1789.539 |  | 2 | 128.6605 | 104.4373 |  | 2 | 243.4565 | 150.1279
Blank | 2.152 | 1.826 | 1.706 | 0 | 0 | Blank | 0 | 0 | 0 | 0 | 0 |  | 3 | 152.6986 | 183.1962 |  | 3 | 280.6833 | 207.6464
0.2 |  | 22.916 | 15.44 | 321.403 | 674.272 | 0.2 |  | 86.257 | 67.636 | 540.326 | 1695.594 |  | 4 | 150.9087 | 197.9199 |  | 4 | 274.3285 | 222.7987
0.2 |  | 22.846 | 15.176 | 324.627 | 672.018 | 0.2 |  | 86.714 | 66.282 | 542.219 | 1715.134 |  | 5 | 169.6512 | 250.5033 |  | 5 | 275.1004 | 227.3759
0.2 |  | 23.122 | 15.205 | 330.18 | 676.595 | 0.2 |  | 84.841 | 0.329 | 533.691 | 1729.01 |  | 6 | 173.2923 | 244.7552 |  | 6 | 283.1312 | 259.4538
Blank |  | 1.779 | 0 | 0 | 0 | Blank |  | 0 | 0 | 0 | 0 |  | 7 | 163.1064 | 243.4622 |  | 7 | 263.3278 | 225.4027
0.25 |  | 17.791 | 11.618 | 289.224 | 457.938 | 0.25 |  | 76.758 | 61.203 | 467.787 | 1219.536 |  | 8 | 201.0123 | 270.6713 |  | 8 | 293.8207 | 263.9126
0.25 |  | 17.683 | 11.79 | 287.637 | 466.979 | 0.25 |  | 76.543 | 0.143 | 457.154 | 1208.746 |  | 12 | 146.0655 | 680.8465 |  | 12 | 240.9252 | 1152.2571
0.25 |  | 17.764 | 11.963 | 289.601 | 466.475 | 0.25 |  | 77.652 | 0.358 | 464.22 | 1201.207 |  | 18 | 190.8948 | 824.549 |  | 18 | 251.0365 | 1285.2022
Blank |  | 1.727 | 1.666 | 0 | 0 | Blank |  | 0 | 0 | 0 | 0 |  | 24 | 142.2005 | 777.9775 |  | 24 | 237.9684 | 1152.4141
0.3 |  | 16.143 | 11.028 | 236.497 | 354.915 | 0.3 |  | 87.674 | 0.046 | 355.751 | 898.652 |  | 48 | 162.2479 | 950.6109 |  | 48 | 261.4561 | 1234.5725
0.3 |  | 16.272 | 10.98 | 217.586 | 387.029 | 0.3 |  | 84.632 | 0.192 | 363.456 | 884.988 | D5 | 120 | 150.7091 | 830.2581 | D5 | 120 | 263.7176 | 1245.329
0.3 |  | 16.015 | 10.908 | 217.51 | 356.268 | 0.3 |  | 85.65 | 0.164 | 358.447 | 875.605 | D8 | 192 | 136.4708 | 750.7659 | D8 | 192 | 237.6892 | 1059.0442
Blank |  | 1.748 | 0 | 0 | 0 | Blank |  | 0 | 0 | 0 | 0 | D11 | 264 | 209.8777 | 1111.4732 | D11 | 264 | 307.6691 | 1570.5051
0.35 |  | 15.612 | 11.834 | 181.375 | 265.903 | 0.35 |  | 81.305 | 0.298 | 227.139 | 613.209 | D14 | 336 | 61.5766 | 706.0528 | D14 | 336 | 84.6661 | 980.7545
0.35 |  | 15.689 | 11.778 | 154.131 | 270.158 | 0.35 |  | 80.11 | 0.209 | 225.881 | 615.287 | D17 | 408 | 54.3988 | 684.7169 | D17 | 408 | 74.2642 | 983.3122
0.35 |  | 15.449 | 11.849 | 179.408 | 272.35 | 0.35 |  | 80.83 | 0.24 | 224.002 | 559.803 | D20 | 480 | 55.1311 | 440.4259 | D20 | 480 | 75.5944 | 630.6473
Blank |  | 1.795 | 1.66 | 0 | 0 | Blank |  | 0 | 0 | 0 | 0 | D23 | 552 | 51.773 | 373.1621 | D23 | 552 | 70.7185 | 539.7809
0.4 |  | 14.951 | 11.59 | 151.567 | 216.962 | 0.4 |  | 80.915 | 0.047 | 143.34 | 356.569 | D26 | 624 | 46.7914 | 220.994 | D26 | 624 | 62.832 | 316.5488
0.4 |  | 15.193 | 11.606 | 121.138 | 216.235 | 0.4 |  | 81.549 | 0.228 | 147.324 | 351.478 | D29 | 696 | 33.9359 | 196.2911 | D29 | 696 | 43.9883 | 285.4475
0.4 |  | 14.522 | 11.473 | 122.942 | 218.037 | 0.4 |  | 82.075 | 0 | 146.565 | 344.737 | D32 | 768 | 28.3352 | 166.7429 | D32 | 768 | 35.0557 | 240.4633
Blank |  | 1.706 | 1.684 | 0 | 0 | Blank |  | 0 | 0 | 0 | 0 | D35 | 840 | 15.2443 | 52.1827 | D35 | 840 | 16.2951 | 69.2937
0.45 |  | 14.28 | 10.837 | 93.618 | 157.515 | 0.45 |  | 76.56 | 0.197 | 61.104 | 213.454 | D38 | 912 | 13.8644 | 28.2229 | D38 | 912 | 14.529 | 33.6507
0.45 |  | 14.173 | 10.852 | 116.992 | 157.395 | 0.45 |  | 76.659 | 0 | 55.891 | 216.869 | D41 | 984 | 6.6731 | 19.5298 | D41 | 984 | 4.1834 | 20.9067
0.45 |  | 14.05 | 10.905 | 93.853 | 159.147 | 0.45 |  | 79.535 | 0.002 | 62.781 | 224.48
Blank |  | 1.739 | 1.66 | 0 | 0 | Blank |  | 0 | 0 | 0 | 0
0.5 |  | 13.269 | 9.524 | 74.16 | 141.621 | 0.5 |  | 74.406 | 0 | 22.681 | 124.13
0.5 |  | 13.564 | 9.542 | 75.828 | 143.583 | 0.5 |  | 74.732 | 0 | 17.07 | 126.156
0.5 |  | 13.138 | 9.535 | 76.256 | 143.836 | 0.5 |  | 72.823 | 0 | 19.891 | 127.094
Blank |  | 1.705 | 0 | 0 | 0 | Blank |  | 0 | 0 | 0 | 0
0.55 |  | 14.027 | 11.721 | 25.933 | 120.951 | 0.55 |  | 83.206 | 0.053 | 0.818 | 69.051
0.55 |  | 14.111 | 11.692 | 26.2 | 120.971 | 0.55 |  | 81.169 | 0.203 | 103.857 | 68.014
0.55 |  | 13.805 | 11.752 | 26.602 | 121.307 | 0.55 |  | 82.522 | 0.116 | 104.468 | 68.723
Blank |  | 0 | 1.675 | 0 | 0 | Blank |  | 0 | 0 | 0 | 0
1 | 993.181 | 14.04 | 11.888 | 20.82 | 104.096 | 1 | 523.114 | 86.208 | 0.221 | 0.489 | 1.445
1 | 992.248 | 14.144 | 11.954 | 19.968 | 103.538 | 1 | 523.573 | 86.809 | 0.101 | 0.431 | 4.407
1 | 996.421 | 13.803 | 11.731 | 20.63 | 104.135 | 1 | 515.531 | 85.155 | 0.107 | 52.142 | 2.125
Blank | 0 | 1.719 | 1.689 | 0 | 0 | Blank | 0 | 0 | 0 | 0 | 0
2 | 597.712 | 15.241 | 12.902 | 8.68 | 7.474 | 2 | 385.597 | 85.607 | 0.221 | 37.26 | 35.395
2 | 597.727 | 15.118 | 12.862 | 8.699 | 7.509 | 2 | 384.712 | 87.056 | 0 | 0.006 | 35.078
2 | 601.645 | 15.101 | 12.817 | 8.847 | 7.574 | 2 | 393.623 | 86.529 | 0.026 | 0.076 | 35.374
Blank | 0 | 1.693 | 1.677 | 0 | 0 | Blank | 0 | 0 | 0 | 0 | 0
3 | 693.179 | 23.779 | 24.397 | 5.509 | 14.699 | 3 | 588.703 | 117.84 | 141.686 | 0 | 77.439
3 | 695.4 | 23.899 | 24.24 | 5.519 | 15.107 | 3 | 558.27 | 116.229 | 165.5 | 0 | 77.018
3 | 695.985 | 23.863 | 24.455 | 5.621 | 14.827 | 3 | 567.717 | 118.613 | 138.526 | 0 | 80.402
Blank | 0 | 0 | 1.703 | 0 | 0 | Blank | 0 | 0 | 0 | 0 | 0
4 | 680.645 | 23.411 | 23.052 | 10.857 | 16.579 | 4 | 629.475 | 119.591 | 130.566 | 0.248 | 93.005
4 | 685.402 | 22.997 | 23.376 | 10.702 | 16.607 | 4 | 617.929 | 117.355 | 129.778 | 0.421 | 92.822
4 | 676.568 | 22.989 | 22.976 | 10.731 | 16.739 | 4 | 612.727 | 118.826 | 153.533 | 60.919 | 91.604
Blank | 0 | 0 | 1.696 | 0 | 0 | Blank | 0 | 0 | 0 | 0 | 0
5 | 698.642 | 57.649 | 62.078 | 12.212 | 17.498 | 5 | 649.524 | 163.155 | 288.118 | 76.268 | 99.483
5 | 700.526 | 56.927 | 62.512 | 11.953 | 17.323 | 5 | 673.735 | 164.343 | 278.884 | 0.498 | 98.466
5 | 699.803 | 56.566 | 61.809 | 12.157 | 17.113 | 5 | 662.072 | 162.286 | 261.305 | 75.588 | 103.825
Blank | 0 | 1.754 | 1.861 | 0 | 0 | Blank | 0 | 0 | 0 | 0 | 0
6 | 718.333 | 59.097 | 65.021 | 13.47 | 11.504 | 6 | 710.764 | 155.847 | 251.336 | 83.39 | 71.208
6 | 720.061 | 57.852 | 64.774 | 13.036 | 11.46 | 6 | 737.515 | 146.765 | 248.269 | 84.071 | 70.642
6 | 717.269 | 58.055 | 65.005 | 12.987 | 11.46 | 6 | 717.973 | 150.085 | 243.231 | 0.232 | 0
Blank | 0 | 1.739 | 1.867 | 0 | 0 | Blank | 0 | 0 | 0 | 0 | 0
7 | 665.745 | 58.678 | 63.771 | 12.786 | 12.168 | 7 | 669.314 | 154.043 | 238.436 | 91.463 | 78.346
7 | 673.63 | 59.112 | 62.06 | 12.652 | 12.198 | 7 | 666.086 | 148.648 | 239.831 | 89.507 | 77.119
7 | 670.887 | 57.15 | 60.832 | 12.785 | 12.142 | 7 | 657.428 | 147.36 | 228.973 | 87.995 | 77.384
Blank | 0 | 1.743 | 1.849 | 0 | 0 | Blank | 0 | 0 | 0 | 0 | 0
8 | 686.482 | 51.842 | 52.089 | 10.831 |  | 8 | 705.624 | 142.532 | 189.709 | 82.761
8 | 691.145 | 56.277 | 51.617 | 11.17 |  | 8 | 696.13 | 142.736 | 188.702 | 83.09
8 | 684.408 | 54.811 | 50.57 | 10.906 |  | 8 | 696.13 | 138.148 | 181.876 | 0.617
Blank | 0 | 1.759 | 1.841 | 0 | 0 | Blank | 0 | 0 | 0 | 0 | 0
12 | 613.185 | 54.03 | 50.649 | 1.522 | 14.381 | 12 | 2915.256 | 140.925 | 184.731 | 0 | 133.092
12 | 609.313 | 54.212 | 50.772 | 1.516 | 14.615 | 12 | 2916.244 | 142.164 | 182.482 | 47.501 | 130.5
12 | 606.918 | 53.777 | 50.211 | 1.494 | 14.387 | 12 | 2882.726 | 138.505 | 184.56 | 47.509 | 166.502
Blank | 0 | 0 | 1.78 | 0 | 0 | Blank | 0 | 0 | 0 | 0 | 0
18 | 597.608 | 129.085 | 33.399 | 1.511 |  | 18 | 2945.827 | 154.109 | 152.938 | 53.573
18 | 599.648 | 130.298 | 33.342 | 1.491 |  | 18 | 2965.419 | 160.922 | 152.172 | 53.836
18 | 599.981 | 129.775 | 33.108 | 1.491 |  | 18 | 2950.882 | 150.583 | 154.327 | 0
Blank | 0 | 0 | 1.714 | 0 |  | Blank | 0 | 0 | 0 | 0
24 | 598.491 | 38.281 | 32.707 | 1.518 | 38.156 | 24 | 2984.885 | 127.199 | 155.272 | 53.86 | 608.289
24 | 604.391 | 38.296 | 32.616 | 1.508 | 37.807 | 24 | 2973.67 | 124.083 | 149.844 | 53.765 | 590.735
24 | 600.707 | 38.112 | 32.702 | 1.501 | 36.214 | 24 | 2947.682 | 125.926 | 150.785 | 0 | 623.667
Blank | 0 | 0 | 1.763 | 0 | 0 | Blank | 0 | 0 | 0 | 0 | 0
48 | 591.764 | 15.938 | 14.405 |  | 23.202 | 48 | 2955.441 | 99.064 | 115.52 |  | 626.518
48 | 598.495 | 16.028 | 14.509 |  | 22.589 | 48 | 2949.507 | 100.53 | 113.717 |  | 623.257
48 | 597.088 | 16.041 | 14.372 |  | 22.544 | 48 | 2990.41 | 97.985 | 116.837 |  | 618.545
Blank | 0 | 1.681 | 1.707 |  | 0 | Blank | 0 | 0 | 0 |  | 0
120 | 588.683 | 4.104 | 3.238 |  | 7.45 | 120 | 2897.95 | 65.793 | 47.449 |  | 319.939
120 | 586.067 | 4.094 | 3.32 |  | 7.5 | 120 | 2905.662 | 77.11 | 47.453 |  | 316.753
120 | 589.326 | 4.109 | 3.282 |  | 7.336 | 120 | 2858.003 | 66.412 | 47.672 |  | 312.901
Blank | 0 | 1.659 | 0 |  | 0 | Blank | 0 | 0 | 0 |  | 0
192 | 527.254 | 3.746 | 3.581 |  | 7.888 | 192 | 2492.422 | 43.677 | 38.562 |  | 451.26
192 | 533.664 | 3.814 | 3.583 |  | 7.905 | 192 | 2524.461 | 43.474 | 38.127 |  | 438.329
192 | 530.93 | 3.795 | 3.589 |  | 7.901 | 192 | 2435.687 | 43.597 | 38.159 |  | 421.436
Blank | 0 | 0 | 0 |  | 0 | Blank | 0 | 0 | 0 |  | 0
264 | 614.522 | 5.044 | 4.512 |  |  | 264 | 3203.022 | 72.494 | 57.695
264 | 622.481 | 5.043 | 4.496 |  |  | 264 | 3224.501 | 70.405 | 56.74
264 | 623.268 | 5.066 | 4.467 |  |  | 264 | 3188.766 | 71.741 | 57.895
Blank | 0 | 0 | 1.667 |  |  | Blank | 0 | 0 | 0
336 | 175.326 | 5.468 | 4.779 |  |  | 336 | 1987.322 | 54.866 | 47.778
336 | 173.865 | 5.44 | 4.855 |  |  | 336 | 2035.427 | 53.608 | 48.844
336 | 174.198 | 5.418 | 4.84 |  |  | 336 | 2018.1 | 59.186 | 49.344
Blank | 2.109 | 0 | 0 |  |  | Blank | 0 | 0 | 0
408 | 156.625 | 5.114 | 4.569 |  |  | 408 | 1980.261 | 33.863 | 27.118
408 | 151.448 | 5.226 | 4.61 |  |  | 408 | 2009.114 | 30.114 | 26.609
408 | 152.125 | 5.22 | 4.652 |  |  | 408 | 1997.907 | 30.26 | 27.206
Blank | 2.088 | 0 | 0 |  |  | Blank | 0 | 0 | 0
480 | 156.516 | 5.06 | 4.415 |  |  | 480 | 1272.589 | 22.699 | 18.608
480 | 154.716 | 5.05 | 4.468 |  |  | 480 | 1300.102 | 20.97 | 18.662
480 | 156.53 | 4.959 | 4.466 |  |  | 480 | 1270.958 | 20.46 | 18.785
Blank | 2.09 | 0 | 0 |  |  | Blank | 0 | 0 | 0
552 | 147.397 | 4.983 | 4.286 |  |  | 552 | 1079.747 | 14.788 | 11.905
552 | 146.405 | 4.958 | 4.303 |  |  | 552 | 1103.471 | 14.842 | 11.993
552 | 144.372 | 4.914 | 4.339 |  |  | 552 | 1095.252 | 14.516 | 11.945
Blank | 2.073 | 0 | 0 |  |  | Blank | 0 | 0 | 0
624 | 132.614 | 5.261 | 4.633 |  |  | 624 | 640.351 | 11.423 | 9.375
624 | 128.93 | 5.249 | 4.611 |  |  | 624 | 629.994 | 11.492 | 8.262
624 | 130.128 | 5.138 | 4.559 |  |  | 624 | 658.493 | 11.42 | 8.136
Blank | 2.067 | 0 | 0 |  |  | Blank | 0 | 0 | 0
696 | 93.767 | 4.886 | 4.339 |  |  | 696 | 584.558 | 6.525 | 5.176
696 | 92.443 | 4.896 | 4.381 |  |  | 696 | 566.688 | 6.551 | 5.353
696 | 91.534 | 4.843 | 4.334 |  |  | 696 | 579.268 | 6.498 | 6.003
Blank | 2.034 | 0 | 0 |  |  | Blank | 0 | 0 | 0
768 | 74.918 | 5.33 | 4.516 |  |  | 768 | 477.525 | 7.438 | 5.715
768 | 74.613 | 5.4 | 4.606 |  |  | 768 | 477.895 | 7.334 | 5.632
768 | 75.686 | 5.373 | 4.575 |  |  | 768 | 506.099 | 7.208 | 5.84
Blank | 2.01 | 0 | 0 |  |  | Blank | 0 | 0 | 0
840 | 37.439 | 4.665 | 4.103 |  |  | 840 | 145.882 | 6.731 | 5.356
840 | 36.974 | 4.671 | 4.101 |  |  | 840 | 144.65 | 6.547 | 5.274
840 | 36.486 | 4.675 | 4.085 |  |  | 840 | 143.167 | 6.752 | 5.285
Blank | 1.966 | 0 | 0 |  |  | Blank | 0 | 0 | 0
912 | 33.64 | 4.464 | 3.89 |  |  | 912 | 73.979 | 6.435 | 5.252
912 | 32.818 | 4.394 | 3.964 |  |  | 912 | 72.27 | 6.366 | 5.131
912 | 33.238 | 4.484 | 3.888 |  |  | 912 | 72.994 | 6.446 | 5.133
Blank | 1.98 | 0 | 0 |  |  | Blank | 0 | 0 | 0
984 | 12.437 | 4.206 | 3.62 |  |  | 984 | 47.305 | 6.165 | 4.836
984 | 12.008 | 4.136 | 3.617 |  |  | 984 | 47.888 | 6.729 | 4.888
984 | 12.275 | 4.15 | 3.609 |  |  | 984 | 46.976 | 6.123 | 4.858
Blank | 0 | 0 | 0 |  |  | Blank | 0 | 0 | 0

## Spacerv+t
| TOB Concentration (µg/mL) |  |  |  |  | VANC Concentration (µg/mL) |  |  |  | Spacer Only (2g Vancomycin & 2g Tobramycin)
[11 empty] | Averages over all repeats |  |  |  | Std Dev over all repeats
| G3A | G3B | G3C | G3D |  | G3A | G3B | G3C | G3D
0 |  | 115.649 |  |  | 0 |  | 14.249
0 |  | 115.169 |  |  | 0 |  | 14.219 |  |  |  | Timepoint | Tob | Vanc |  | Timepoint | Tob | Vanc
0 |  | 114.533 |  |  | 0 |  | 14.004 |  |  |  | 0 | 115.117 | 14.1573 |  | 0 | 0.55981 | 0.13364
Blank |  | 0 |  |  | Blank |  | 0 |  |  |  | 0.05 | 72.0122 | 41.1212 |  | 0.05 | 33.2283 | 22.3523
5 | 38.816 | 123.672 | 68.845 | 56.968 | 5 | 9.895 | 36.074 | 50.965 | 64.516 |  | 0.1 | 64.0064 | 32.9189 |  | 0.1 | 35.4387 | 15.7432
5 | 38.597 | 123.917 | 68.298 | 57.001 | 5 | 10.175 | 35.979 | 50.462 | 63.211 |  | 0.15 | 58.6692 | 34.6628 |  | 0.15 | 39.1664 | 19.6748
5 | 39.078 | 124.447 | 67.904 | 56.603 | 5 | 9.554 | 35.946 | 51.48 | 75.198 |  | 0.2 | 52.5211 | 22.5764 |  | 0.2 | 39.4267 | 11.0223
Blank | 1.194 | 0 | 0 | 0 | Blank | 0 | 0 | 0 | 0 |  | 0.25 | 45.9605 | 21.5148 |  | 0.25 | 39.9055 | 12.1757
10 | 31.531 | 118.187 | 66.079 | 40.969 | 10 | 7.522 | 38.117 | 40.937 | 46.293 |  | 0.3 | 16.898 | 14.7793 |  | 0.3 | 6.898 | 7.8794
10 | 31.737 | 118.384 | 65.705 | 39.788 | 10 | 7.09 | 37.304 | 40.314 | 44.471 |  | 0.35 | 16.2187 | 12.7625 |  | 0.35 | 9.9553 | 6.0145
10 | 30.505 | 118.768 | 66.4 | 40.024 | 10 | 7.292 | 37.932 | 41.174 | 46.581 |  | 0.4 | 13.7854 | 11.8777 |  | 0.4 | 9.3674 | 8.5361
Blank | 1.183 | 0 | 0 | 0 | Blank | 0 | 0 | 0 | 0 |  | 0.45 | 11.9603 | 8.973 |  | 0.45 | 7.3182 | 4.0387
15 | 17.937 | 115.005 | 69.097 | 33.454 | 15 | 7.488 | 28.227 | 57.552 | 42.081 |  | 0.5 | 10.3963 | 9.0453 |  | 0.5 | 6.8633 | 5.81
15 | 17.865 | 115.259 | 68.276 | 33.35 | 15 | 7.296 | 28.16 | 58.684 | 46.853 |  | 0.55 | 9.5254 | 7.5372 |  | 0.55 | 5.675 | 4.9299
15 | 17.645 | 116.035 | 66.991 | 33.117 | 15 | 7.621 | 28.816 | 56.279 | 46.897 |  | 1 | 10.7764 | 7.1704 |  | 1 | 9.3311 | 5.8876
Blank | 1.175 | 0 | 0 | 0 | Blank | 0 | 0 | 0 | 0 |  | 2 | 4.7742 | 1.4571 |  | 2 | 2.5352 | 1.4162
20 | 17.379 | 112.094 | 57.187 | 23.321 | 20 | 5.781 | 21.606 | 30.257 | 31.847 |  | 3 | 4.0885 | 1.3989 |  | 3 | 2.0698 | 0.95603
20 | 17.365 | 112.12 | 58.643 | 22.646 | 20 | 5.764 | 20.927 | 30.749 | 33.601 |  | 4 | 5.3854 | 2.6094 |  | 4 | 3.9821 | 2.0951
20 | 17.051 | 112.227 | 57.312 | 22.908 | 20 | 5.913 | 21.653 | 31.123 | 31.696 |  | 5 | 5.8426 | 1.8401 |  | 5 | 3.6253 | 1.8586
Blank | 1.179 | 0 | 0 | 0 | Blank | 0 | 0 | 0 | 0 |  | 6 | 5.9012 | 1.718 |  | 6 | 4.6966 | 2.6449
25 | 12.205 | 107.979 | 46.066 | 17.616 | 25 | 5.081 | 17.444 | 36.823 | 27.872 |  | 7 | 6.9638 | 1.5175 |  | 7 | 4.284 | 2.2319
25 | 11.912 | 108.34 | 46.877 | 17.206 | 25 | 5.042 | 17.137 | 35.546 | 27.418 |  | 8 | 6.9622 | 7.6413 |  | 8 | 3.076 | 8.7395
25 | 12.019 | 108.173 | 45.934 | 17.199 | 25 | 4.801 | 17.145 | 35.903 | 27.966 |  | 12 | 5.0438 | 2.3763 |  | 12 | 1.6773 | 2.2766
Blank | 1.177 | 0 | 0 | 0 | Blank | 0 | 0 | 0 | 0 |  | 18 | 8.2193 | 0.69644 |  | 18 | 5.1267 | 0.61394
30 | 11.584 | 12.476 | 28.518 | 15.512 | 30 | 5.444 | 12.34 | 14.801 | 26.524 |  | 24 | 11.3953 | 6.0207 |  | 24 | 8.4046 | 10.3624
30 | 11.602 | 12.616 | 28.242 | 15.382 | 30 | 5.461 | 12.634 | 15.007 | 26.134 |  | 48 | 7.6107 | 4.7905 |  | 48 | 5.266 | 5.2482
30 | 11.569 | 12.611 | 27.464 | 15.2 | 30 | 5.223 | 12.192 | 15.269 | 26.322 | D5 | 120 | 9.0333 | 4.3839 | D5 | 120 | 7.4349 | 5.0404
Blank | 0 | 0 | 0 | 0 | Blank | 0 | 0 | 0 | 0 | D8 | 192 | 9.1263 | 6.2632 | D8 | 192 | 5.4778 | 4.348
35 | 8.688 | 12.353 | 32.88 | 11.67 | 35 | 3.007 | 15.495 | 17.441 | 15.034 | D11 | 264 | 8.832 | 4.5797 | D11 | 264 | 8.084 | 4.6764
35 | 8.575 | 11.949 | 32.496 | 11.533 | 35 | 2.841 | 15.777 | 17.735 | 15.199 | D14 | 336 | 8.7782 | 1.6567 | D14 | 336 | 6.9601 | 1.6689
35 | 8.624 | 12.122 | 32.317 | 11.417 | 35 | 2.87 | 15.073 | 17.133 | 15.545 | D17 | 408 | 5.4265 | 2.2293 | D17 | 408 | 4.5146 | 2.871
Blank | 0 | 0 | 0 | 0 | Blank | 0 | 0 | 0 | 0 | D20 | 480 | 3.9517 | 2.674 | D20 | 480 | 2.1159 | 2.0817
40 | 6.694 | 10.508 | 29.192 | 8.975 | 40 | 1.086 | 9.233 | 23.592 | 13.699 | D23 | 552 | 4.1607 | 2.4535 | D23 | 552 | 2.2605 | 1.5334
40 | 6.568 | 10.387 | 29.045 | 8.799 | 40 | 1.166 | 8.506 | 23.45 | 13.701 | D26 | 624 | 3.3812 | 2.5402 | D26 | 624 | 2.0079 | 1.0333
40 | 6.573 | 10.392 | 29.179 | 9.113 | 40 | 1.324 | 8.878 | 24.135 | 13.762 | D29 | 696 | 3.1987 | 1.2153 | D29 | 696 | 1.3881 | 0.96928
Blank | 0 | 0 | 0 | 0 | Blank | 0 | 0 | 0 | 0 | D32 | 768 | 3.1788 | 1.0265 | D32 | 768 | 1.4741 | 1.1246
45 | 6.529 | 10.127 | 23.765 | 7.62 | 45 | 2.818 | 9.591 | 13.176 | 10.272 | D35 | 840 | 3.694 | 2.5683 | D35 | 840 | 1.8678 | 2.8135
45 | 6.391 | 10.026 | 23.933 | 7.519 | 45 | 2.845 | 9.535 | 13.463 | 10.289 | D38 | 912 | 3.4357 | 3.044 | D38 | 912 | 2.1485 | 2.5393
45 | 6.245 | 9.888 | 23.947 | 7.533 | 45 | 2.609 | 9.419 | 13.414 | 10.245 | D41 | 984 | 3.5755 | 3.6795 | D41 | 984 | 1.7725 | 3.4154
Blank | 0 | 0 | 0 | 0 | Blank | 0 | 0 | 0 | 0
50 | 4.243 | 10.377 | 20.878 | 5.876 | 50 | 0.366 | 10.253 | 15.982 | 9.934
50 | 4.271 | 10.386 | 21.133 | 5.85 | 50 | 0.356 | 9.706 | 15.615 | 10.008
50 | 4.203 | 10.391 | 21.254 | 5.893 | 50 | 0.272 | 10.187 | 15.857 | 10.007
Blank | 0 | 0 | 0 | 0 | Blank | 0 | 0 | 0 | 0
55 | 4.027 | 11.797 | 17.308 | 4.906 | 55 | 0.718 | 12.861 | 10.886 | 5.89
55 | 4.078 | 11.816 | 17.442 | 4.898 | 55 | 0.647 | 12.591 | 10.866 | 5.663
55 | 4.024 | 11.732 | 17.368 | 4.909 | 55 | 0.711 | 13.034 | 10.841 | 5.738
Blank | 0 | 0 | 0 | 0 | Blank | 0 | 0 | 0 | 0
60 | 3.834 | 9.221 | 25.891 | 4.136 | 60 | 0.2 | 9.435 | 15.176 | 3.811
60 | 3.905 | 9.419 | 25.852 | 4.118 | 60 | 0.296 | 9.674 | 15.088 | 3.912
60 | 3.847 | 9.406 | 25.59 | 4.098 | 60 | 0.253 | 9.344 | 15.023 | 3.833
Blank | 0 | 0 | 0 | 0 | Blank | 0 | 0 | 0 | 0
2 | 2.94 | 3.838 | 9.038 | 3.364 | 2 | 0 | 0.695 | 3.523 | 1.629
2 | 3.064 | 3.775 | 9.019 | 3.363 | 2 | 0 | 0.665 | 3.55 | 1.57
2 | 3.028 | 3.688 | 8.796 | 3.377 | 2 | 0 | 0.632 | 3.718 | 1.503
Blank | 0 | 0 | 0 | 0 | Blank | 0 | 0 | 0 | 0
3 | 1.83 | 3.281 | 7.291 | 4.044 | 3 | 0.598 | 0.799 | 1.463 | 3.002
3 | 1.826 | 3.291 | 7.214 | 3.961 | 3 | 0.531 | 0.772 | 1.327 | 2.85
3 | 1.826 | 3.246 | 7.21 | 4.042 | 3 | 0.54 | 0.872 | 1.167 | 2.866
BLANK | 0 | 0 | 0 | 0 | BLANK | 0 | 0 | 0 | 0
4 | 1.967 | 3.236 | 11.784 | 4.593 | 4 | 0.603 | 0.799 | 5.301 | 3.638
4 | 1.925 | 3.278 | 11.937 | 4.56 | 4 | 0.614 | 0.842 | 5.457 | 3.477
4 | 1.922 | 3.26 | 11.651 | 4.512 | 4 | 0.594 | 0.836 | 5.517 | 3.635
BLANK | 0 | 0 | 0 | 0 | BLANK | 0 | 0 | 0 | 0
5 | 4.099 | 2.795 | 11.731 | 4.749 | 5 | 0.437 | 0 | 3.848 | 3.48
5 | 4.021 | 2.789 | 11.735 | 4.834 | 5 | 0 | 0 | 3.819 | 3.325
5 | 4.124 | 2.76 | 11.706 | 4.768 | 5 | 0 | 0 | 3.871 | 3.301
BLANK | 0 | 0 | 0 | 0 | BLANK | 0 | 0 | 0 | 0
6 | 4.326 | 2.453 | 13.809 | 3.346 | 6 | 0 | 0 | 5.939 | 0.83
6 | 4.138 | 2.438 | 13.702 | 3.337 | 6 | 0 | 0 | 6.205 | 0.763
6 | 4.148 | 2.436 | 13.325 | 3.356 | 6 | 0 | 0 | 6.053 | 0.826
BLANK | 0 | 0 | 0 | 0 | BLANK | 0 | 0 | 0 | 0
7 | 10.727 | 2.447 | 11.502 | 3.327 | 7 | 0 | 0 | 4.977 | 0.915
7 | 10.576 | 2.456 | 11.46 | 3.312 | 7 | 0 | 0 | 5.277 | 0.892
7 | 10.517 | 2.46 | 11.471 | 3.31 | 7 | 0 | 0 | 5.228 | 0.921
BLANK | 0 | 0 | 0 | 0 | BLANK | 0 | 0 | 0 | 0
8 | 4.759 | 4.954 | 11.112 |  | 8 | 0 | 19.133 | 3.88
8 | 4.873 | 5.031 | 11.035 |  | 8 | 0 | 19.04 | 3.83
8 | 4.772 | 5.088 | 11.036 |  | 8 | 0 | 19.069 | 3.82
BLANK | 1.178 | 0 | 0 | 0 | BLANK | 0 | 0 | 0 | 0
12 | 6.456 | 3.279 | 5.851 | 3.938 | 12 | 0 | 6.016 | 1.468 | 2.118
12 | 6.384 | 3.3 | 5.843 | 3.892 | 12 | 0 | 5.993 | 1.421 | 2.066
12 | 8.521 | 3.297 | 5.874 | 3.891 | 12 | 0 | 5.716 | 1.58 | 2.138
BLANK | 1.194 | 0 | 0 | 0 | BLANK | 0 | 0 | 0 | 0
18 | 15.06 | 3.669 | 6.068 |  | 18 | 0 | 0.675 | 1.445
18 | 14.908 | 3.695 | 6.079 |  | 18 | 0 | 0.739 | 1.466
18 | 14.774 | 3.673 | 6.048 |  | 18 | 0 | 0.615 | 1.328
BLANK | 1.202 | 0 | 0 | 0 | BLANK | 0 | 0 | 0 | 0
24 | 13.77 | 3.528 | 5.589 | 23.95 | 24 | 0 | 0.886 | 0 | 23.285
24 | 11.287 | 3.551 | 5.604 | 24.364 | 24 | 0 | 0.886 | 0 | 23.045
24 | 11.933 | 3.532 | 5.616 | 24.019 | 24 | 0 | 0.897 | 0 | 23.249
BLANK | 0 | 0 | 0 | 0 | BLANK | 0 | 0 | 0 | 0
48 | 2.809 |  |  | 12.625 | 48 | 0 |  |  | 9.579
48 | 2.793 |  |  | 12.361 | 48 | 0 |  |  | 9.689
48 | 2.812 |  |  | 12.264 | 48 | 0 |  |  | 9.475
BLANK | 0 |  |  | 0 | BLANK | 0 |  |  | 0
D5 | 4.443 | 18.627 |  | 3.781 | D5 | 0 | 11.052 |  | 2.23
D5 | 4.316 | 19.275 |  | 3.83 | D5 | 0 | 11.043 |  | 2.095
D5 | 4.333 | 18.914 |  | 3.781 | D5 | 0 | 10.865 |  | 2.17
BLANK | 0 | 0 |  | 0 | BLANK | 0 | 0 |  | 0
D8 | 3.642 | 16.224 |  | 7.672 | D8 | 0.745 | 10.509 |  | 7.424
D8 | 3.64 | 15.89 |  | 7.691 | D8 | 0.754 | 10.571 |  | 7.509
D8 | 3.62 | 16.008 |  | 7.75 | D8 | 0.685 | 10.483 |  | 7.689
BLANK | 0 | 0 |  | 0 | BLANK | 0 | 0 |  | 0
D11 | 1.454 | 16.295 |  |  | D11 | 0.338 | 9.28
D11 | 1.448 | 16.252 |  |  | D11 | 0.303 | 8.843
D11 | 1.456 | 16.087 |  |  | D11 | 0.314 | 8.4
BLANK | 0 | 0 |  |  | BLANK | 0 | 0
D14 | 2.45 | 15.316 |  |  | D14 | 0.421 | 3.119
D14 | 2.423 | 15.108 |  |  | D14 | 0 | 3.263
D14 | 2.403 | 14.969 |  |  | D14 | 0 | 3.137
BLANK | 0 | 0 |  |  | BLANK | 0 | 0
D17 | 1.316 | 9.665 |  |  | D17 | 0.529 | 0
D17 | 1.298 | 9.483 |  |  | D17 | 0.513 | 5.854
D17 | 1.303 | 9.494 |  |  | D17 | 0.481 | 5.999
BLANK | 0 | 0 |  |  | BLANK | 0 | 0
D20 | 1.997 | 5.924 |  |  | D20 | 1.414 | 4.655
D20 | 2.04 | 5.927 |  |  | D20 | 0.51 | 4.549
D20 | 2.025 | 5.797 |  |  | D20 | 0.475 | 4.441
BLANK | 0 | 0 |  |  | BLANK | 0 | 0
D23 | 2.117 | 6.284 |  |  | D23 | 1.043 | 3.902
D23 | 2.082 | 6.182 |  |  | D23 | 1.109 | 3.853
D23 | 2.093 | 6.206 |  |  | D23 | 1.011 | 3.803
BLANK | 0 | 0 |  |  | BLANK | 0 | 0
D26 | 1.542 | 5.119 |  |  | D26 | 1.585 | 3.406
D26 | 1.554 | 5.22 |  |  | D26 | 1.618 | 3.479
D26 | 1.551 | 5.301 |  |  | D26 | 1.591 | 3.562
BLANK | 0 | 0 |  |  | BLANK | 0 | 0
D29 | 1.949 | 4.461 |  |  | D29 | 0.344 | 2.105
D29 | 1.922 | 4.497 |  |  | D29 | 0.315 | 2.122
D29 | 1.924 | 4.439 |  |  | D29 | 0.333 | 2.073
BLANK | 0 | 0 |  |  | BLANK | 0 | 0
D32 | 1.845 | 4.546 |  |  | D32 | 0 | 2.085
D32 | 1.829 | 4.542 |  |  | D32 | 0 | 2.045
D32 | 1.826 | 4.485 |  |  | D32 | 0 | 2.029
BLANK | 0 | 0 |  |  | BLANK | 0 | 0
D35 | 2.003 | 5.38 |  |  | D35 | 0 | 5.161
D35 | 1.984 | 5.431 |  |  | D35 | 0 | 5.127
D35 | 1.98 | 5.386 |  |  | D35 | 0 | 5.122
BLANK | 0 | 0 |  |  | BLANK | 0 | 0
D38 | 1.472 | 5.393 |  |  | D38 | 0.752 | 5.331
D38 | 1.476 | 5.456 |  |  | D38 | 0.716 | 5.377
D38 | 1.476 | 5.341 |  |  | D38 | 0.71 | 5.378
BLANK | 0 | 0 |  |  | BLANK | 0 | 0
D41 | 1.961 | 5.109 |  |  | D41 | 0.908 | 6.596
D41 | 1.976 | 5.232 |  |  | D41 | 0.822 | 6.823
D41 | 1.937 | 5.238 |  |  | D41 | 0 | 6.928
BLANK | 0 | 0 |  |  | BLANK | 0 | 0

## VP
| TOB Concentraion (µg/mL) |  |  |  | VANC Concentration (µg/mL) |  |  | 2g Vancomycin Dump with an Unloaded Spacer
[9 empty] | Averages over all repeats |  |  |  | Std Dev over all repeats
| VD1A | VD1B | VD1C |  | VD1A | VD1B | VD1C
0 | 0 | 0 |  | 0 | 0 | 8176.146
0 | 0 | 0 |  | 0 | 0 | 8062.697 |  |  | Timepoint | Tob | Vanc |  | Timepoint | Tob | Vanc
0 | 0 | 0 |  | 0 | 0 | 8345.544 |  |  | 0 | 0 | 4097.3978 |  | 0 | 0 | 4489.3772
Blank | 0 | 0 |  | Blank | 0 | 0 | 0 |  | 0.05 | 0 | 5075.3461 |  | 0.05 | 0 | 3808.6856
5 | 0 | 0 | 0 | 5 | 7628.782 | 7802.841 | 0 |  | 0.1 | 0 | 3210.4466 |  | 0.1 | 0 | 2428.5746
5 | 0 | 0 | 0 | 5 | 7811.975 | 7461.587 | 0 |  | 0.15 | 0 | 2068.0558 |  | 0.15 | 0 | 1568.4432
5 | 0 | 0 | 0 | 5 | 7515.944 | 7456.986 | 0 |  | 0.2 | 0 | 1457.0153 |  | 0.2 | 0 | 944.2293
Blank | 0 | 0 | 0 | Blank | 0 | 0 | 0 |  | 0.25 | 0 | 769.4243 |  | 0.25 | 0 | 577.6838
10 | 0 | 0 | 0 | 10 | 5198.52 | 4453.828 | 0 |  | 0.3 | 0 | 384.0398 |  | 0.3 | 0 | 323.7716
10 | 0 | 0 | 0 | 10 | 5147.306 | 4455.158 | 0 |  | 0.35 | 0 | 142.7196 |  | 0.35 | 0 | 200.1266
10 | 0 | 0 | 0 | 10 | 5197.067 | 4442.14 | 0 |  | 0.4 | 0 | 166.1273 |  | 0.4 | 0 | 32.4925
Blank | 0 | 0 | 0 | Blank | 0 | 0 | 0 |  | 0.45 | 0 | 87.4868 |  | 0.45 | 0 | 66.4772
15 | 0 | 0 | 0 | 15 | 3333.629 | 2831.324 | 0 |  | 0.5 | 0 | 49.4746 |  | 0.5 | 0 | 37.9083
15 | 0 | 0 | 0 | 15 | 3365.401 | 2881.125 | 0 |  | 0.55 | 0 | 18.4124 |  | 0.55 | 0 | 27.625
15 | 0 | 0 | 0 | 15 | 3408.065 | 2792.958 | 0 |  | 1 | 0 | 12.3689 |  | 1 | 0 | 16.6934
Blank | 0 | 0 | 0 | Blank | 0 | 0 | 0 |  | 2 | 0 | 0 |  | 2 | 0 | 0
20 | 0 | 0 | 0 | 20 | 2374.526 | 1811.193 | 236.451 |  | 3 | 0 | 0 |  | 3 | 0 | 0
20 | 0 | 0 | 0 | 20 | 2382.916 | 1774.804 | 245.185 |  | 4 | 0 | 0 |  | 4 | 0 | 0
20 | 0 | 0 | 0 | 20 | 2253.312 | 1800.969 | 233.782 |  | 5 | 0 | 0 |  | 5 | 0 | 0
Blank | 0 | 0 | 0 | Blank | 0 | 0 | 0 |  | 6 | 0 | 0 |  | 6 | 0 | 0
25 | 0 | 0 | 0 | 25 | 1335.373 | 923.957 | 38.14 |  | 7 | 0 | 0 |  | 7 | 0 | 0
25 | 0 | 0 | 0 | 25 | 1348.087 | 942.098 | 36.987 |  | 8 | 0 | 0 |  | 8 | 0 | 0
25 | 0 | 0 | 0 | 25 | 1345.818 | 916.849 | 37.51 |  | 12 | 0 | 0 |  | 12 | 0 | 0
Blank | 0 | 0 | 0 | Blank | 0 | 0 | 0 |  | 18 | 0 | 0 |  | 18 | 0 | 0
30 | 0 | 0 | 0 | 30 | 753.796 | 402.567 | 5.098 |  | 24 | 0 | 0 |  | 24 | 0 | 0
30 | 0 | 0 | 0 | 30 | 748.017 | 391.356 | 4.503 |  | 48 | 0 | 0 |  | 48 | 0 | 0
30 | 0 | 0 | 0 | 30 | 755.255 | 390.593 | 5.173 | D5 | 120 | 0 | 0 | D5 | 120 | 0 | 0
Blank | 0 | 0 | 0 | Blank | 0 | 0 | 0 | D8 | 192 | 0 | 0 | D8 | 192 | 0 | 0
35 | 0 | 0 | 0 | 35 | 416.569 | 19.222 | 0.097 | D11 | 264 | 0 | 0 | D11 | 264 | 0 | 0
35 | 0 | 0 | 0 | 35 | 403.321 | 20.085 | 0 | D14 | 336 | 0 | 0 | D14 | 336 | 0 | 0
35 | 0 | 0 | 0 | 35 | 407.994 | 17.188 | 0 | D17 | 408 | 0 | 0 | D17 | 408 | 0 | 0
Blank | 0 | 0 | 0 | Blank | 0 | 0 | 0 | D20 | 480 | 0 | 0 | D20 | 480 | 0 | 0
40 | 0 | 0 |  | 40 | 193.282 | 138.165 |  | D23 | 552 | 0 | 0 | D23 | 552 | 0 | 0
40 | 0 | 0 |  | 40 | 198.144 | 134.804 |  | D26 | 624 | 0 | 0 | D26 | 624 | 0 | 0
40 | 0 | 0 |  | 40 | 195.793 | 136.576 |  | D29 | 696 | 0 | 0 | D29 | 696 | 0 | 0
Blank | 0 | 0 |  | Blank | 0 | 0 |  | D32 | 768 | #DIV/0! | #DIV/0! | D32 | 768 | #DIV/0! | #DIV/0!
45 | 0 | 0 | 0 | 45 | 142.245 | 117.845 | 0 | D35 | 840 | #DIV/0! | #DIV/0! | D35 | 840 | #DIV/0! | #DIV/0!
45 | 0 | 0 | 0 | 45 | 142.971 | 117.468 | 0 | D38 | 912 | #DIV/0! | #DIV/0! | D38 | 912 | #DIV/0! | #DIV/0!
45 | 0 | 0 | 0 | 45 | 145.124 | 121.728 | 0 | D41 | 984 | #DIV/0! | #DIV/0! | D41 | 984 | #DIV/0! | #DIV/0!
Blank | 0 | 0 | 0 | Blank | 0 | 0 | 0
50 | 0 | 0 | 0 | 50 | 63.303 | 82.379 | 0
50 | 0 | 0 | 0 | 50 | 67.764 | 81.942 | 0
50 | 0 | 0 | 0 | 50 | 65.107 | 84.776 | 0
Blank | 0 | 0 | 0 | Blank | 0 | 0 | 0
55 | 0 | 0 | 0 | 55 | 0 | 54.434 | 0
55 | 0 | 0 | 0 | 55 | 0 | 56.598 | 0
55 | 0 | 0 | 0 | 55 | 0 | 54.68 | 0
Blank | 0 | 0 | 0 | Blank | 0 | 0 | 0
60 | 0 | 0 | 0 | 60 | 0 | 34.741 | 2.449
60 | 0 | 0 | 0 | 60 | 0 | 34.354 | 2.583
60 | 0 | 0 | 0 | 60 | 0 | 34.64 | 2.553
Blank | 0 | 0 | 0 | Blank | 0 | 0 | 0
2 | 0 | 0 | 0 | 2 | 0 | 0 | 0
2 | 0 | 0 | 0 | 2 | 0 | 0 | 0
2 | 0 | 0 | 0 | 2 | 0 | 0 | 0
Blank | 0 | 0 | 0 | Blank | 0 | 0 | 0
3 | 0 | 0 | 0 | 3 | 0 | 0 | 0
3 | 0 | 0 | 0 | 3 | 0 | 0 | 0
3 | 0 | 0 | 0 | 3 | 0 | 0 | 0
BLANK | 0 | 0 | 0 | BLANK | 0 | 0 | 0
4 | 0 | 0 | 0 | 4 | 0 | 0 | 0
4 | 0 | 0 | 0 | 4 | 0 | 0 | 0
4 | 0 | 0 | 0 | 4 | 0 | 0 | 0
BLANK | 0 | 0 | 0 | BLANK | 0 | 0 | 0
5 | 0 | 0 | 0 | 5 | 0 | 0 | 0
5 | 0 | 0 | 0 | 5 | 0 | 0 | 0
5 | 0 | 0 | 0 | 5 | 0 | 0 | 0
BLANK | 0 | 0 | 0 | BLANK | 0 | 0 | 0
6 | 0 | 0 | 0 | 6 | 0 | 0 | 0
6 | 0 | 0 | 0 | 6 | 0 | 0 | 0
6 | 0 | 0 | 0 | 6 | 0 | 0 | 0
BLANK | 0 | 0 | 0 | BLANK | 0 | 0 | 0
7 | 0 | 0 |  | 7 | 0 | 0
7 | 0 | 0 |  | 7 | 0 | 0
7 | 0 | 0 |  | 7 | 0 | 0
BLANK | 0 | 0 |  | BLANK | 0 | 0
8 | 0 | 0 | 0 | 8 | 0 | 0 | 0
8 | 0 | 0 | 0 | 8 | 0 | 0 | 0
8 | 0 | 0 | 0 | 8 | 0 | 0 | 0
BLANK | 0 | 0 | 0 | BLANK | 0 | 0 | 0
12 | 0 | 0 | 0 | 12 | 0 | 0 | 0
12 | 0 | 0 | 0 | 12 | 0 | 0 | 0
12 | 0 | 0 | 0 | 12 | 0 | 0 | 0
BLANK | 0 | 0 | 0 | BLANK | 0 | 0 | 0
18 | 0 | 0 | 0 | 18 | 0 | 0 | 0
18 | 0 | 0 | 0 | 18 | 0 | 0 | 0
18 | 0 | 0 | 0 | 18 | 0 | 0 | 0
BLANK | 0 | 0 | 0 | BLANK | 0 | 0 | 0
24 | 0 | 0 |  | 24 | 0 | 0
24 | 0 | 0 |  | 24 | 0 | 0
24 | 0 | 0 |  | 24 | 0 | 0
BLANK | 0 | 0 |  | BLANK | 0 | 0
48 |  |  | 0 | 48 |  |  | 0
48 |  |  | 0 | 48 |  |  | 0
48 |  |  | 0 | 48 |  |  | 0
BLANK |  |  | 0 | BLANK | 0 |  | 0
D5 | 0 | 0 | 0 | D5 | 0 | 0 | 0
D5 | 0 | 0 | 0 | D5 | 0 | 0 | 0
D5 | 0 | 0 | 0 | D5 | 0 | 0 | 0
BLANK | 0 | 0 | 0 | BLANK | 0 | 0 | 0
D8 | 0 | 0 |  | D8 | 0 | 0
D8 | 0 | 0 |  | D8 | 0 | 0
D8 | 0 | 0 |  | D8 | 0 | 0
BLANK | 0 | 0 |  | BLANK | 0 | 0
D11 | 0 | 0 |  | D11 | 0 | 0
D11 | 0 | 0 |  | D11 | 0 | 0
D11 | 0 | 0 |  | D11 | 0 | 0
BLANK | 0 | 0 |  | BLANK | 0 | 0
D14 | 0 | 0 |  | D14 | 0 | 0
D14 | 0 | 0 |  | D14 | 0 | 0
D14 | 0 | 0 |  | D14 | 0 | 0
BLANK | 0 | 0 |  | BLANK | 0 | 0
D17 | 0 | 0 |  | D17 | 0 | 0
D17 | 0 | 0 |  | D17 | 0 | 0
D17 | 0 | 0 |  | D17 | 0 | 0
BLANK | 0 | 0 |  | BLANK | 0 | 0
D20 | 0 | 0 |  | D20 | 0 | 0
D20 | 0 | 0 |  | D20 | 0 | 0
D20 | 0 | 0 |  | D20 | 0 | 0
BLANK | 0 | 0 |  | BLANK | 0 | 0
D23 | 0 | 0 |  | D23 | 0 | 0
D23 | 0 | 0 |  | D23 | 0 | 0
D23 | 0 | 0 |  | D23 | 0 | 0
BLANK | 0 | 0 |  | BLANK | 0 | 0
D26 | 0 | 0 |  | D26 | 0 | 0
D26 | 0 | 0 |  | D26 | 0 | 0
D26 | 0 | 0 |  | D26 | 0 | 0
BLANK | 0 | 0 |  | BLANK | 0 | 0
D29 | 0 | 0 |  | D29 | 0 | 0
D29 | 0 | 0 |  | D29 | 0 | 0
D29 | 0 | 0 |  | D29 | 0 | 0
BLANK | 0 | 0 |  | BLANK | 0 | 0
D32 |  |  |  | D32
D32 |  |  |  | D32
D32 |  |  |  | D32
BLANK |  |  |  | BLANK
D35 |  |  |  | D35
D35 |  |  |  | D35
D35 |  |  |  | D35
BLANK |  |  |  | BLANK
D38 |  |  |  | D38
D38 |  |  |  | D38
D38 |  |  |  | D38
BLANK |  |  |  | BLANK
D41 |  |  |  | D41
D41 |  |  |  | D41
D41 |  |  |  | D41
BLANK |  |  |  | BLANK

## Vancomycin
VANC-UAMS1 |  | CSBv+tSpaceru data was not used because it was clinically irrelevant | [56 empty] | GEOMETRIC MEAN | Standard Error
| Log-scale | CSB+Spaceru |  |  |  | CSBv+tSpacerv+t |  |  |  | Spacerv+t |  |  |  |  | VP |  |  | Geometric mean | CSB+Spaceru |  |  |  | CSBv+tSpacerv+t |  |  |  | Spacerv+t |  |  |  |  | VP |  | Time (Days) | Time (Hr) | Average |  |  |  |  | Average |  |  |  |  | Average | [10 empty] | Flow Rate
|  | G1B | G1C | G1E | G1F | G2B | G2C | G2D | G2E | G3A | G3B | G3C | G3D | [5 empty] | G1B | G1C | G1E | G1F | G2B | G2C | G2D | G2E | G3A | G3B | G3C | G3D | VP1 | VP2 | VP3 |  |  | CSBv+tplusSpaceru | SD | SE | CI | n | CSBv+tplusSpacerv+t | SD | SE | CI | n | Spacerv+t | SD | SE | CI | n | VP | SD | SE | CI | n | Min. Detection | Hour | mL/min
time (min) | time (hr) |  |  |  |  | CSB+S.ii | CSB+S.iii | CSB+S.iv | CSB+S.v | S.i | S.ii | S.iii | S.iv | VP.i | VP.ii | VP.iii | time (min) | time (hr) |  |  |  |  | CSB+S.ii | CSB+S.iii | CSB+S.iv | CSB+S.v | S.i | S.ii | S.iii | S.iv | VP.i | VP.ii | VP1.iii | 0.0026042 | 0.0625 | 2.0675 | 1.4335 | 0.71677 | 1.3977 | 4 | 2.536 | 0.56242 | 0.28121 | 0.54836 | 4 | 1.5217 | 0.36882 | 0.18441 | 0.36144 | 4 | 3.8815 | 0.0031943 | 0.0022587 | 0.0044045 | 2 | -0.01 | 0.0625 | 3.5
5 | 0.083 | -0.0087739 | 2.3208 | 2.6424 | 3.2447 | 2.0631 | 2.1914 | 2.8606 | 3.1704 | 0.99542 | 1.5572 | 1.7073 | 1.8097 | 3.8825 | 3.8923 | -0.01 | 5 | 0.083 | -0.0087739 | 2.3106 | 2.7449 | 3.2235 | 2.0602 | 2.0632 | 2.8539 | 3.1668 | 0.99438 | 1.5563 | 1.7073 | 1.8289 | 3.8837 | 3.8792 |  | 0.0069583 | 0.167 | 2.2392 | 0.77377 | 0.38688 | 0.75442 | 4 | 2.334 | 0.85838 | 0.42919 | 0.83692 | 4 | 1.428 | 0.37802 | 0.18901 | 0.37046 | 4 | 3.6814 | 0.046676 | 0.033005 | 0.064359 | 2 | -0.01 | 1 | 3.5
5 | 0.083 | -0.0087739 | 2.3152 | 2.7981 | 3.2403 | 2.0594 | 1.998 | 2.8621 | 3.173 | 1.0075 | 1.556 | 1.703 | 1.8008 | 3.8928 | 3.8728 | -0.01 | 10" | 0.167 | 1.2503 | 2.0786 | 2.5571 | 3.0707 | 2.0196 | 1.2795 | 2.8416 | 3.1953 | 0.86328 | 1.5773 | 1.6107 | 1.6606 | 3.7144 | 3.6484 |  | 0.010417 | 0.25 | 2.0523 | 0.89085 | 0.44542 | 0.86858 | 4 | 2.4694 | 0.66974 | 0.33487 | 0.653 | 4 | 1.4354 | 0.39579 | 0.19789 | 0.38787 | 4 | 3.49 | 0.052997 | 0.037474 | 0.073075 | 2 | -0.01 | 1 | 3.13
5 | 0.083 | -0.0087739 | 2.2956 | 2.7942 | 3.1855 | 2.058 | 2.0001 | 2.8391 | 3.157 | 0.98019 | 1.5557 | 1.7116 | 1.8762 | 3.876 | 3.8726 | -0.01 | 15" | 0.25 | 1.1559 | 1.4269 | 2.6939 | 2.9324 | 1.9605 | 1.8625 | 2.8084 | 3.2462 | 0.87315 | 1.4533 | 1.7596 | 1.6553 | 3.5275 | 3.4525 |  | 0.013875 | 0.333 | 1.9029 | 0.81054 | 0.40527 | 0.79027 | 4 | 2.2784 | 0.88861 | 0.4443 | 0.86639 | 4 | 1.2731 | 0.34816 | 0.17408 | 0.34119 | 4 | 3 | 0.54227 | 0.31308 | 0.6105 | 3 | -0.01 | 1.9999 | 3.13
10 | 0.167 | 1.2499 | 2.0749 | 2.5507 | 3.065 | 2.0204 | 1.9226 | 2.8386 | 3.1891 | 0.87633 | 1.5811 | 1.6121 | 1.6655 | 3.7159 | 3.6487 | -0.01 | 20" | 0.333 | 1.154 | 1.2761 | 2.4027 | 2.7789 | 1.9342 | 1.2143 | 2.7314 | 3.2338 | 0.76484 | 1.3303 | 1.4872 | 1.5101 | 3.3685 | 3.2542 | 2.3773 | 0.017375 | 0.417 | 1.7784 | 0.76236 | 0.38118 | 0.7433 | 4 | 2.0561 | 1.0961 | 0.54806 | 1.0687 | 4 | 1.2335 | 0.38172 | 0.19086 | 0.37408 | 4 | 2.5567 | 0.85434 | 0.49325 | 0.96185 | 3 | -0.01 | 2 | 1.54
10 | 0.167 | 1.2455 | 2.0775 | 2.5512 | 3.0672 | 2.0217 | 1.9246 | 2.8391 | 3.1974 | 0.85065 | 1.5718 | 1.6055 | 1.6481 | 3.7116 | 3.6489 | -0.01 | 25" | 0.417 | 1.0994 | 1.1737 | 2.2045 | 2.6359 | 1.8864 | 0.58974 | 2.6656 | 3.0827 | 0.69663 | 1.2366 | 1.5573 | 1.4433 | 3.1281 | 2.9673 | 1.5745 | 0.020833 | 0.5 | 1.6277 | 0.64812 | 0.32406 | 0.63192 | 4 | 1.8571 | 1.312 | 0.65601 | 1.2792 | 4 | 1.1051 | 0.28584 | 0.14292 | 0.28012 | 4 | 2.0548 | 1.1889 | 0.68639 | 1.3385 | 3 | -0.01 | 3 | 1.54
10 | 0.167 | 1.2557 | 2.0834 | 2.5693 | 3.08 | 2.0167 | -0.0087739 | 2.8473 | 3.1995 | 0.86285 | 1.579 | 1.6146 | 1.6682 | 3.7158 | 3.6476 | -0.01 | 30" | 0.5 | 1.1044 | 1.2002 | 1.6837 | 2.5224 | 1.9344 | -0.0087739 | 2.5553 | 2.9476 | 0.73037 | 1.093 | 1.1768 | 1.4204 | 2.8764 | 2.5964 | 0.69155 | 0.024292 | 0.583 | 1.4693 | 0.61832 | 0.30916 | 0.60286 | 4 | 1.7567 | 1.2291 | 0.61457 | 1.1984 | 4 | 1.0192 | 0.37162 | 0.18581 | 0.36419 | 4 | 1.2924 | 1.3105 | 0.75661 | 1.4754 | 3 | -0.01 | 3 | 1.02
| 0.25 | 1.1731 | 1.4256 | 2.6816 | 2.933 | 1.9757 | 1.8683 | 2.8245 | 3.2384 | 0.87437 | 1.4507 | 1.7601 | 1.6241 | 3.5229 | 3.452 | -0.01 | 35" | 0.583 | 1.1321 | 1.2296 | 1.1216 | 2.3939 | 1.9071 | -0.0087739 | 2.3535 | 2.7749 | 0.46316 | 1.1888 | 1.2414 | 1.1835 | 2.612 | 1.274 | -0.0087739 | 0.027792 | 0.667 | 1.41 | 0.65624 | 0.32812 | 0.63983 | 4 | 1.6535 | 1.1365 | 0.56827 | 1.1081 | 4 | 0.88379 | 0.56696 | 0.28348 | 0.55562 | 4 | 2.2134 | 0.11066 | 0.078249 | 0.15259 | 2 | -0.01 | 4 | 1.02
| 0.25 | 1.1505 | 1.4276 | 2.7062 | 2.9315 | 1.9524 | 1.8626 | 2.808 | 3.2475 | 0.86308 | 1.4496 | 1.7685 | 1.6707 | 3.527 | 3.4596 | -0.01 | 40" | 0.667 | 1.1331 | 1.2223 | 0.91013 | 2.3744 | 1.9112 | -0.0058493 | 2.1636 | 2.5452 | 0.074805 | 0.94779 | 1.3752 | 1.1374 | 2.2917 | 2.1352 |  | 0.03125 | 0.75 | 1.3588 | 0.57236 | 0.28618 | 0.55805 | 4 | 1.4992 | 1.0342 | 0.51711 | 1.0084 | 4 | 0.88889 | 0.30571 | 0.15286 | 0.2996 | 4 | 1.4078 | 1.2275 | 0.70868 | 1.3819 | 3 | -0.01 | 4 | 0.76
| 0.25 | 1.1442 | 1.4276 | 2.6938 | 2.9328 | 1.9535 | 1.8565 | 2.7927 | 3.2527 | 0.88201 | 1.4596 | 1.7503 | 1.6711 | 3.5325 | 3.4461 | -0.01 | 45" | 0.75 | 1.1409 | 1.1891 | 0.90779 | 2.1973 | 1.8897 | -0.0087739 | 1.7771 | 2.3389 | 0.44017 | 0.9784 | 1.1255 | 1.0115 | 2.1567 | 2.0755 | -0.0087739 | 0.034708 | 0.833 | 1.3324 | 0.53572 | 0.26786 | 0.52233 | 4 | 1.3139 | 0.94436 | 0.47218 | 0.92075 | 4 | 0.7979 | 0.54586 | 0.27293 | 0.53495 | 4 | 1.2419 | 1.0844 | 0.62607 | 1.2208 | 3 | -0.01 | 5 | 0.76
| 0.333 | 1.1557 | 1.2804 | 2.41 | 2.7702 | 1.9358 | 1.8302 | 2.7327 | 3.2293 | 0.762 | 1.3346 | 1.4808 | 1.5031 | 3.3756 | 3.258 | 2.3737 | 50" | 0.833 | 1.1281 | 1.1886 | 0.89896 | 2.1139 | 1.8691 | -0.0087739 | 1.2955 | 2.0996 | -0.0087739 | 1.002 | 1.1991 | 0.99926 | 1.8154 | 1.9192 | -0.0087739 | 0.038208 | 0.917 | 1.5282 | 0.45092 | 0.22546 | 0.43965 | 4 | 1.2713 | 0.89025 | 0.44513 | 0.868 | 4 | 0.724 | 0.51095 | 0.25547 | 0.50073 | 4 | 0.57487 | 1.0109 | 0.58365 | 1.1381 | 3 | -0.01 | 5 | 0.61
| 0.333 | 1.152 | 1.2773 | 2.4013 | 2.7848 | 1.9381 | 1.8214 | 2.7342 | 3.2343 | 0.76072 | 1.3207 | 1.4878 | 1.5264 | 3.3771 | 3.2492 | 2.3895 | 55" | 0.917 | 1.1252 | 1.2004 | 1.702 | 2.0849 | 1.9154 | -0.0087739 | 1.3422 | 1.8363 | -0.0087739 | 1.1081 | 1.036 | 0.76064 | -0.0087739 | 1.7422 | -0.0087739 | 0.041667 | 1 | 1.7547 | 0.38762 | 0.19381 | 0.37793 | 4 | 0.71742 | 0.84614 | 0.42307 | 0.82499 | 4 | 0.68317 | 0.52291 | 0.26146 | 0.51245 | 4 | 0.64425 | 0.80156 | 0.46278 | 0.90242 | 3 | -0.01 | 6 | 0.61
| 0.333 | 1.1544 | 1.2706 | 2.3967 | 2.7818 | 1.9286 | -0.0087739 | 2.7273 | 3.2378 | 0.77181 | 1.3355 | 1.4931 | 1.501 | 3.3528 | 3.2555 | 2.3688 | 60" | 1 | 1.2088 | 2.0435 | 1.7489 | 2.0179 | 1.9348 | -0.0087739 | 0.56655 | 0.37712 | -0.0087739 | 0.97696 | 1.1788 | 0.58566 | -0.0087739 | 1.5388 | 0.40272 | 0.083333 | 2 | 1.3299 | 0.31371 | 0.15685 | 0.30587 | 4 | 0.99829 | 0.89942 | 0.44971 | 0.87693 | 4 | 0.1833 | 0.26626 | 0.13313 | 0.26094 | 4 | -0.0087739 | 0 | 0 | 0 | 3 | -0.01 | 6 | 0.5
| 0.417 | 1.0983 | 1.1787 | 2.2077 | 2.6221 | 1.8851 | 1.7868 | 2.67 | 3.0862 | 0.70595 | 1.2416 | 1.5661 | 1.4452 | 3.1256 | 2.9657 | 1.5814 | 2' | 2 | 1.1932 | 1.2867 | 1.7794 | 1.0603 | 1.9365 | -0.0087739 | 0.5179 | 1.5476 | -0.0087739 | -0.0087739 | 0.55582 | 0.19493 | -0.0087739 | -0.0087739 | -0.0087739 | 0.125 | 3 | 1.4471 | 0.2132 | 0.1066 | 0.20787 | 4 | 1.5314 | 1.0332 | 0.51658 | 1.0073 | 4 | 0.14101 | 0.22299 | 0.11149 | 0.21853 | 4 | -0.0087739 | 0 | 0 | 0 | 3 | -0.01 | 7 | 0.5
| 0.417 | 1.1025 | 1.1687 | 2.1875 | 2.6347 | 1.8839 | -0.0087739 | 2.6601 | 3.0823 | 0.7026 | 1.2339 | 1.5508 | 1.438 | 3.1297 | 2.9741 | 1.568 | 3' | 3 | 1.4327 | 1.4439 | 1.7166 | 1.1951 | 2.0702 | 2.1706 | -0.0087739 | 1.8936 | -0.0087739 | -0.0087739 | 0.1184 | 0.46318 | -0.0087739 | -0.0087739 | -0.0087739 | 0.16667 | 4 | 1.6314 | 0.36758 | 0.18379 | 0.35839 | 4 | 1.6919 | 0.73864 | 0.36932 | 0.72018 | 4 | 0.31775 | 0.38414 | 0.19207 | 0.37646 | 4 | -0.0087739 | 0 | 0 | 0 | 3 | -0.01 | 7 | 0.43
| 0.417 | 1.0974 | 1.1738 | 2.2183 | 2.651 | 1.8902 | -0.0087739 | 2.6667 | 3.0796 | 0.68133 | 1.2341 | 1.5551 | 1.4466 | 3.129 | 2.9623 | 1.5741 | 4' | 4 | 2.1125 | 1.4744 | 1.6892 | 1.2495 | 2.074 | 2.1384 | 0.58907 | 1.966 | -0.0087739 | -0.0087739 | 0.73434 | 0.55419 | -0.0087739 | -0.0087739 | -0.0087739 | 0.20833 | 5 | 1.5981 | 0.41101 | 0.2055 | 0.40073 | 4 | 1.9767 | 0.51603 | 0.25802 | 0.50313 | 4 | 0.2737 | 0.32702 | 0.16351 | 0.32048 | 4 | -0.0087739 | 0 | 0 | 0 | 3 | -0.01 | 8 | 0.43
| 0.5 | 1.1041 | 1.1811 | 1.8891 | 2.5338 | 1.9429 | -0.0087739 | 2.5511 | 2.9536 | 0.73592 | 1.0913 | 1.1703 | 1.4236 | 2.8773 | 2.6048 | 0.7074 | 5' | 5 | 1.4881 | 2.1981 | 1.4391 | 1.2673 | 2.2129 | 2.4407 | 1.2507 | 2.0024 | -0.0087739 | -0.0087739 | 0.585 | 0.52734 | -0.0087739 | -0.0087739 | -0.0087739 | 0.25 | 6 | 1.7798 | 0.55467 | 0.27733 | 0.5408 | 4 | 1.7706 | 0.60212 | 0.30106 | 0.58706 | 4 | 0.18912 | 0.39579 | 0.1979 | 0.38788 | 4 | -0.0087739 | 0 | 0 | 0 | 3 | -0.01 | 8 | 0.38
| 0.5 | 1.1071 | 1.1845 | 1.664 | 2.5316 | 1.9275 | -0.0087739 | 2.5605 | 2.9469 | 0.73727 | 1.1015 | 1.1763 | 1.4172 | 2.8739 | 2.5926 | 0.6535 | 6' | 6 | 2.2344 | 2.2041 | 1.6142 | 1.0665 | 2.1786 | 2.3937 | 1.279 | 1.2309 | -0.0087739 | -0.0087739 | 0.78281 | -0.0087739 | -0.0087739 | -0.0087739 | -0.0087739 | 0.29167 | 7 | 1.7143 | 0.43273 | 0.21636 | 0.42191 | 4 | 2.0977 | 0.22047 | 0.11023 | 0.21495 | 4 | 0.17156 | 0.36067 | 0.18033 | 0.35346 | 4 | -0.0087739 | 0 | 0 | 0 | 2 | -0.01 | 12 | 0.38
| 0.5 | 1.1019 | 1.2351 | 1.4978 | 2.5017 | 1.9327 | -0.0087739 | 2.5544 | 2.9423 | 0.71792 | 1.0861 | 1.1838 | 1.4203 | 2.8781 | 2.5917 | 0.71374 | 7' | 7 | 1.6991 | 2.1352 | 1.9001 | 1.1228 | 2.1761 | 2.3724 | 1.9525 | 1.8899 | -0.0087739 | -0.0087739 | 0.71256 | -0.0087739 | -0.0087739 | -0.0087739 |  | 0.33333 | 8 | 2.0821 | 0.088213 | 0.05093 | 0.099313 | 3 | 1.899 | 0.54278 | 0.31337 | 0.61108 | 3 | 0.61884 | 0.64536 | 0.3726 | 0.73029 | 3 | -0.0087739 | 0 | 0 | 0 | 3 | -0.01 | 12 | 0.25
| 0.583 | 1.129 | 1.2283 | -0.0087739 | 2.3891 | 1.9101 | -0.0087739 | 2.3563 | 2.7876 | 0.47813 | 1.1902 | 1.2416 | 1.1771 | 2.6197 | 1.2838 | -0.0087739 | 8' | 8 | 2.151 | 2.1128 | 1.9827 |  | 2.1496 | 2.2712 | 1.2762 |  | -0.0087739 | 1.2806 | 0.5847 |  | -0.0087739 | -0.0087739 | -0.0087739 | 0.5 | 12 | 1.8565 | 0.4273 | 0.21365 | 0.41661 | 4 | 1.9203 | 0.53958 | 0.26979 | 0.52609 | 4 | 0.31473 | 0.33338 | 0.16669 | 0.32671 | 4 | -0.0087739 | 0 | 0 | 0 | 3 | -0.01 | 18 | 0.25
| 0.583 | 1.1353 | 1.2261 | 1.8452 | 2.398 | 1.9037 | -0.0087739 | 2.3539 | 2.7891 | 0.45347 | 1.198 | 1.2488 | 1.1818 | 2.6057 | 1.3029 | -0.0087739 | 12' | 12 | 2.1062 | 2.0681 | 2.0345 | 1.2171 | 2.1477 | 2.2646 | 1.1149 | 2.1537 | -0.0087739 | 0.77135 | 0.17266 | 0.32369 | -0.0087739 | -0.0087739 | -0.0087739 | 0.75 | 18 | 2.0698 | 0.035602 | 0.020555 | 0.040082 | 3 | 1.8421 | 0.59901 | 0.34584 | 0.67438 | 3 | 0.044062 | 0.091514 | 0.052836 | 0.10356 | 3 | -0.0087739 | 0 | 0 | 0 | 3 | -0.01 | 18 | 0.16
| 0.583 | 1.132 | 1.2344 | 1.5284 | 2.3946 | 1.9076 | -0.0087739 | 2.3503 | 2.748 | 0.45788 | 1.1782 | 1.2338 | 1.1916 | 2.6107 | 1.2352 | -0.0087739 | 18' | 18 | 2.0522 | 2.0464 | 2.1108 |  | 2.1907 | 2.1851 | 1.1504 |  | -0.0087739 | -0.0087739 | 0.14973 |  | -0.0087739 | -0.0087739 | -0.0087739 | 1 | 24 | 2.0785 | 0.088356 | 0.044178 | 0.086147 | 4 | 2.0539 | 0.67477 | 0.33738 | 0.6579 | 4 | 0.33476 | 0.68706 | 0.34353 | 0.67332 | 4 | -0.0087739 | 0 | 0 | 0 | 2 | -0.01 | 24 | 0.16
| 0.667 | 1.1353 | 1.2216 | 0.91651 | 2.3667 | 1.908 | -0.0087739 | 2.1564 | 2.5521 | 0.03583 | 0.96534 | 1.3728 | 1.1367 | 2.2862 | 2.1404 |  | 24' | 24 | 2.0465 | 2.0287 | 2.2104 | 2.0282 | 2.0994 | 2.1817 | 1.151 | 2.7835 | -0.0087739 | -0.0087739 | -0.0087739 | 1.3654 | -0.0087739 | -0.0087739 |  | 2 | 48 | 1.5202 | 0.8437 | 0.48711 | 0.94987 | 3 | 2.2843 | 0.44294 | 0.25573 | 0.49867 | 3 | 0.23877 | 0.49508 | 0.24754 | 0.48518 | 4 | -0.0087739 | 0 | 0 | 0 | 1 | -0.01 | 24 | 0.12
| 0.667 | 1.1343 | 1.2256 | 0.90913 | 2.379 | 1.9114 | -0.0087739 | 2.1683 | 2.5459 | 0.066699 | 0.92973 | 1.3701 | 1.1368 | 2.297 | 2.1297 |  | 48' | 48 | 0.55429 | 2.113 |  | 1.8934 | 1.9965 | 2.062 |  | 2.7943 | -0.0087739 | -0.0087739 | -0.0087739 | 0.98139 |  |  | -0.0087739 | [23 empty] | 48 | 0.12
| 0.667 | 1.1297 | 1.2197 | 0.90477 | 2.3777 | 1.9142 | 0 | 2.166 | 2.5375 | 0.12189 | 0.94832 | 1.3826 | 1.1387 | 2.2918 | 2.1354 | [41 empty] | 48 | 0.016
| 0.75 | 1.1392 | 1.1887 | 0.90972 | 2.1947 | 1.884 | -0.0087739 | 1.7861 | 2.3293 | 0.44994 | 0.98186 | 1.1198 | 1.0117 | 2.153 | 2.0713 | -0.0087739 |  | Mean | CSBv+tSpacerv+t |  |  |  | Spacerv+t |  |  |  |  | VP | [5 empty] | log difference | [22 empty] | 120 | 0.016
| 0.75 | 1.1452 | 1.1914 | 0.90714 | 2.198 | 1.8846 | -0.0087739 | 1.7473 | 2.3362 | 0.45408 | 0.97932 | 1.1291 | 1.0124 | 2.1552 | 2.0699 | -0.0087739 |  |  | G2B | G2C | G2D | G2E | G3A | G3B | G3C | G3D | VP1 | VP2 | VP3 | [5 empty] | Time (Days) | Time (Hr) | Average | Average | [18 empty] | 120 | 0.016
| 0.75 | 1.1384 | 1.1872 | 0.9065 | 2.1993 | 1.9006 | -0.0087739 | 1.7978 | 2.3512 | 0.41647 | 0.974 | 1.1276 | 1.0105 | 2.1617 | 2.0854 | -0.0087739 |  | time (hour) | CSB+S.ii | CSB+S.iii | CSB+S.iv | CSB+S.v | S.i | S.ii | S.iii | S.iv | VP.i | VP.ii | VP1.iii | [7 empty] | CSBv+tplusSpacerv+t | Spacerv+t | Fold | [17 empty] | 192 | 0.016
| 0.833 | 1.1268 | 1.1888 | 0.90091 | 2.1104 | 1.8716 | -0.0087739 | 1.3557 | 2.0939 | -0.0087739 | 1.0109 | 1.2036 | 0.99712 | 1.8014 | 1.9158 | -0.0087739 |  | 0.083 | 114.8597 | 118.3143 | 714.5773 | 1468.3737 | 9.8747 | 35.9997 | 50.969 | 67.6417 | 7652.2337 | 7573.8047 | [6 empty] | 0.0026042 | 0.0625 | 2.536 | 1.5217 | 1.0143 | [17 empty] | 264 | 0.016
| 0.833 | 1.1289 | 1.1888 | 0.89944 | 2.1142 | 1.8735 | -0.0087739 | 1.2322 | 2.1009 | -0.0087739 | 0.98704 | 1.1935 | 1.0003 | 1.831 | 1.9135 | -0.0087739 |  | 0.167 | 104.6143 | 56.2383 | 694.4817 | 1567.982 | 7.3013 | 37.7843 | 40.8083 | 45.7817 | 5180.9643 | 4450.3753 | [6 empty] | 0.0069583 | 0.167 | 2.334 | 1.428 | 0.90603 | [17 empty] | 336 | 0.016
| 0.833 | 1.1286 | 1.1883 | 0.89653 | 2.1169 | 1.8623 | -0.0087739 | 1.2987 | 2.1041 | -0.0087739 | 1.008 | 1.2002 | 1.0003 | 1.8136 | 1.9283 | -0.0087739 |  | 0.25 | 91.3403 | 72.8593 | 643.575 | 1762.958 | 7.4683 | 28.401 | 57.505 | 45.277 | 3369.0317 | 2835.1357 | [6 empty] | 0.010417 | 0.25 | 2.4694 | 1.4354 | 1.034 | [17 empty] | 408 | 0.016
| 0.917 | 1.1128 | 1.2028 | 1.7029 | 2.0782 | 1.9202 | -0.0087739 | -0.0087739 | 1.8392 | -0.0087739 | 1.1093 | 1.0369 | 0.77012 | -0.0087739 | 1.7359 | -0.0087739 |  | 0.333 | 85.9373 | 44.966 | 538.7453 | 1713.246 | 5.8193 | 21.3953 | 30.7097 | 32.3813 | 2336.918 | 1795.6553 | 238.4727 | [5 empty] | 0.013875 | 0.333 | 2.2784 | 1.2731 | 1.0053 | [17 empty] | 480 | 0.016
| 0.917 | 1.133 | 1.2129 | 1.705 | 2.0863 | 1.9094 | -0.0087739 | 2.0164 | 1.8326 | -0.0087739 | 1.1001 | 1.0361 | 0.75305 | -0.0087739 | 1.7528 | -0.0087739 |  | 0.417 | 76.9843 | 21.0543 | 463.0537 | 1209.8297 | 4.9747 | 17.242 | 36.0907 | 27.752 | 1343.0927 | 927.6347 | 37.5457 | [5 empty] | 0.017375 | 0.417 | 2.0561 | 1.2335 | 0.82266 | [17 empty] | 552 | 0.016
| 0.917 | 1.1299 | 1.1856 | 1.6982 | 2.0902 | 1.9166 | -0.0087739 | 2.019 | 1.8371 | -0.0087739 | 1.1151 | 1.0351 | 0.75876 | -0.0087739 | 1.7378 | -0.0087739 |  | 0.5 | 85.9853 | 0.98 | 359.218 | 886.415 | 5.376 | 12.3887 | 15.0257 | 26.3267 | 752.356 | 394.8387 | 4.9247 | [5 empty] | 0.020833 | 0.5 | 1.8571 | 1.1051 | 0.75201 | [17 empty] | 624 | 0.016
| 1 | 1.2096 | 2.0451 | 1.7439 | 2.0174 | 1.9355 | -0.0087739 | -0.0087739 | 0.15987 | -0.0087739 | 0.97474 | 1.1812 | 0.58104 | -0.0087739 | 1.5408 | 0.38899 |  | 0.583 | 80.7483 | 0.98 | 225.674 | 596.0997 | 2.906 | 15.4483 | 17.4363 | 15.2593 | 409.2947 | 18.8317 | 0.98 | [5 empty] | 0.024292 | 0.583 | 1.7567 | 1.0192 | 0.73747 | [17 empty] | 696 | 0.016
| 1 | 1.2073 | 2.0424 | 1.7504 | 2.0152 | 1.9386 | -0.0087739 | -0.0087739 | 0.64414 | -0.0087739 | 0.98561 | 1.1786 | 0.5924 | -0.0087739 | 1.536 | 0.41212 |  | 0.667 | 81.513 | 0.98 | 145.743 | 350.928 | 1.192 | 8.8723 | 23.7257 | 13.7207 | 195.7397 | 136.515 | [6 empty] | 0.027792 | 0.667 | 1.6535 | 0.88379 | 0.76974 | [17 empty] | 768 | 0.016
| 1 | 1.2093 | 2.0429 | 1.7522 | 2.0209 | 1.9302 | -0.0087739 | 1.7172 | 0.32736 | -0.0087739 | 0.97053 | 1.1768 | 0.58354 | -0.0087739 | 1.5396 | 0.40705 |  | 0.75 | 77.5847 | 0.98 | 59.9253 | 218.2677 | 2.7573 | 9.515 | 13.351 | 10.2687 | 143.4467 | 119.0137 | 0.98 | [5 empty] | 0.03125 | 0.75 | 1.4992 | 0.88889 | 0.61033 | [17 empty] | 840 | 0.016
| 2 | 1.1879 | 1.2829 | 1.7808 | 1.0547 | 1.9325 | -0.0087739 | 1.5712 | 1.5489 | -0.0087739 | -0.0087739 | 0.54691 | 0.21192 | -0.0087739 | -0.0087739 | e |  | 0.833 | 73.987 | 0.98 | 19.8807 | 125.7933 | 0.98 | 10.0487 | 15.818 | 9.983 | 65.3913 | 83.0323 | 0.98 | [5 empty] | 0.034708 | 0.833 | 1.3139 | 0.7979 | 0.51598 | [17 empty] | 912 | 0.016
| 2 | 1.1972 | 1.2871 | 1.7802 | 1.064 | 1.9398 | -0.0087739 | -0.0087739 | 1.545 | -0.0087739 | -0.0087739 | 0.55023 | 0.1959 | -0.0087739 | -0.0087739 | -0.0087739 |  | 0.917 | 82.299 | 0.98 | 69.7683 | 68.596 | 0.98 | 12.8287 | 10.8643 | 5.7637 | 0.98 | 55.2373 | 0.98 | [5 empty] | 0.038208 | 0.917 | 1.2713 | 0.724 | 0.54727 | [17 empty] | 984 | 0.016
| 2 | 1.1944 | 1.2899 | 1.7771 | 1.0622 | 1.9372 | -0.0087739 | -0.0087739 | 1.5487 | -0.0087739 | -0.0087739 | 0.57031 | 0.17696 | -0.0087739 | -0.0087739 | -0.0087739 |  | 1 | 86.0573 | 0.98 | 18.034 | 2.659 | 0.98 | 9.4843 | 15.0957 | 3.852 | 0.98 | 34.5783 | 2.5283 | [5 empty] | 0.041667 | 1 | 0.71742 | 0.68317 | 0.034244
| 3 | 1.4332 | 1.4439 | 1.7102 | 1.1976 | 2.0713 | 2.1513 | -0.0087739 | 1.889 | -0.0087739 | -0.0087739 | 0.16524 | 0.47741 | -0.0087739 | -0.0087739 | -0.0087739 |  | 2 | 86.3973 | 0.98 | 13.0733 | 35.2823 | 0.98 | 0.98 | 3.597 | 1.5673 | 0.98 | 0.98 | 0.98 | [5 empty] | 0.083333 | 2 | 0.99829 | 0.1833 | 0.81499
| 3 | 1.4369 | 1.4461 | 1.7227 | 1.1932 | 2.0653 | 2.2188 | -0.0087739 | 1.8866 | -0.0087739 | -0.0087739 | 0.12287 | 0.45484 | -0.0087739 | -0.0087739 | -0.0087739 |  | 3 | 117.5607 | 148.5707 | 0.98 | 78.2863 | 0.98 | 0.98 | 1.319 | 2.906 | 0.98 | 0.98 | 0.98 | [5 empty] | 0.125 | 3 | 1.5314 | 0.14101 | 1.3904
| 3 | 1.4279 | 1.4417 | 1.7168 | 1.1944 | 2.0741 | 2.1415 | -0.0087739 | 1.9053 | -0.0087739 | -0.0087739 | 0.067071 | 0.45728 | -0.0087739 | -0.0087739 | -0.0087739 |  | 4 | 118.5907 | 137.959 | 20.9597 | 92.477 | 0.98 | 0.98 | 5.425 | 3.5833 | 0.98 | 0.98 | 0.98 | [5 empty] | 0.16667 | 4 | 1.6919 | 0.31775 | 1.3741
| 4 | 2.1111 | 1.4758 | 1.6828 | 1.2526 | 2.0777 | 2.1158 | -0.0087739 | 1.9685 | -0.0087739 | -0.0087739 | 0.72436 | 0.56086 | -0.0087739 | -0.0087739 | -0.0087739 |  | 5 | 163.2613 | 276.1023 | 50.9453 | 100.5913 | 0.98 | 0.98 | 3.846 | 3.3687 | 0.98 | 0.98 | 0.98 | [5 empty] | 0.20833 | 5 | 1.9767 | 0.2737 | 1.703
| 4 | 2.1128 | 1.4723 | 1.6911 | 1.2468 | 2.0695 | 2.1132 | -0.0087739 | 1.9677 | -0.0087739 | -0.0087739 | 0.73695 | 0.5412 | -0.0087739 | -0.0087739 | -0.0087739 |  | 6 | 150.899 | 247.612 | 56.147 | 47.61 | 0.98 | 0.98 | 6.0657 | 0.98 | 0.98 | 0.98 | 0.98 | [5 empty] | 0.25 | 6 | 1.7706 | 0.18912 | 1.5814
| 4 | 2.1137 | 1.475 | 1.6937 | 1.2491 | 2.0749 | 2.1862 | 1.7848 | 1.9619 | -0.0087739 | -0.0087739 | 0.7417 | 0.5605 | -0.0087739 | -0.0087739 | -0.0087739 |  | 7 | 150.017 | 235.7467 | 89.655 | 77.6163 | 0.98 | 0.98 | 5.1607 | 0.98 | 0.98 | 0.98 | [6 empty] | 0.29167 | 7 | 2.0977 | 0.17156 | 1.9262
| 5 | -0.0087739 | 2.1911 | 1.4427 | 1.27 | 2.2126 | 2.4596 | 1.8823 | 1.9977 | -0.0087739 | -0.0087739 | 0.58524 | 0.54158 | -0.0087739 | -0.0087739 | -0.0087739 |  | 8 | 141.1387 | 186.7623 | 55.6103 |  | 0.98 | 19.0807 | 3.8433 |  | 0.98 | 0.98 | 0.98 | [5 empty] | 0.33333 | 8 | 1.899 | 0.61884 | 1.2802
| 5 | 2.2381 | 2.2084 | 1.4386 | 1.2664 | 2.2158 | 2.4454 | -0.0087739 | 1.9933 | -0.0087739 | -0.0087739 | 0.58195 | 0.52179 | -0.0087739 | -0.0087739 | -0.0087739 |  | 12 | 140.5313 | 183.9243 | 31.9967 | 143.3647 | 0.98 | 5.9083 | 1.4897 | 2.1073 | 0.98 | 0.98 | 0.98 | [5 empty] | 0.5 | 12 | 1.9203 | 0.31473 | 1.6055
| 5 | 2.235 | 2.1946 | 1.436 | 1.2654 | 2.2103 | 2.4171 | 1.8785 | 2.0163 | -0.0087739 | -0.0087739 | 0.58782 | 0.51865 | -0.0087739 | -0.0087739 | -0.0087739 |  | 18 | 155.2047 | 153.1457 | 36.1297 |  | 0.98 | 0.98 | 1.413 |  | 0.98 | 0.98 | 0.98 | [5 empty] | 0.75 | 18 | 1.8421 | 0.044062 | 1.798
| 6 | 2.2347 | 2.2006 | 1.62 | 1.0697 | 2.1927 | 2.4003 | 1.9211 | 1.8525 | -0.0087739 | -0.0087739 | 0.77371 | -0.0087739 | -0.0087739 | -0.0087739 | -0.0087739 |  | 24 | 125.736 | 151.967 | 36.2017 | 607.5637 | 0.98 | 0.98 | 0.98 | 23.193 | 0.98 | 0.98 | [6 empty] | 1 | 24 | 2.0539 | 0.33476 | 1.7191
| 6 | 2.2362 | 2.207 | 1.6096 | 1.0651 | 2.1666 | 2.3949 | 1.9246 | 1.8491 | -0.0087739 | -0.0087739 | 0.79274 | -0.0087739 | -0.0087739 | -0.0087739 | -0.0087739 |  | 48 | 99.193 | 115.358 |  | 622.7733 | 0.98 |  |  | 9.581 |  |  | 0.98 | [5 empty] | 2 | 48 | 2.2843 | 0.23877 | 2.0455
| 6 | 2.2322 | 2.2048 | 1.6131 | 1.0646 | 2.1763 | 2.386 | -0.0087739 | -0.0087739 | -0.0087739 | -0.0087739 | 0.78197 | -0.0087739 | -0.0087739 | -0.0087739 | -0.0087739
| 7 | 1.701 | 2.132 | 1.8973 | 1.1221 | 2.1876 | 2.3774 | 1.9612 | 1.894 | -0.0087739 | -0.0087739 | 0.69697 | -0.0087739 | -0.0087739 | -0.0087739
| 7 | 1.702 | 2.1334 | 1.9019 | 1.1229 | 2.1722 | 2.3799 | 1.9519 | 1.8872 | -0.0087739 | -0.0087739 | 0.72239 | -0.0087739 | -0.0087739 | -0.0087739
| 7 | 1.6942 | 2.1401 | 1.9009 | 1.1234 | 2.1684 | 2.3598 | 1.9445 | 1.8887 | -0.0087739 | -0.0087739 | 0.71834 | -0.0087739 | -0.0087739 | -0.0087739
| 8 | 2.1453 | 2.1058 | 1.9876 |  | 2.1539 | 2.2781 | 1.9178 |  | -0.0087739 | 1.2818 | 0.58883 |  | -0.0087739 | -0.0087739 | -0.0087739
| 8 | 2.1526 | 2.1176 | 1.9837 |  | 2.1545 | 2.2758 | 1.9195 |  | -0.0087739 | 1.2797 | 0.5832 |  | -0.0087739 | -0.0087739 | -0.0087739
| 8 | 2.1549 | 2.1148 | 1.9768 |  | 2.1403 | 2.2598 | -0.0087739 |  | -0.0087739 | 1.2803 | 0.58206 |  | -0.0087739 | -0.0087739 | -0.0087739
| 12 | 2.1033 | 2.0695 | 2.0353 | 1.2187 | 2.149 | 2.2665 | -0.0087739 | 2.1242 | -0.0087739 | 0.77931 | 0.16673 | 0.32593 | -0.0087739 | -0.0087739 | -0.0087739
| 12 | 2.1073 | 2.0664 | 2.031 | 1.2169 | 2.1528 | 2.2612 | 1.6767 | 2.1156 | -0.0087739 | 0.77764 | 0.15259 | 0.31513 | -0.0087739 | -0.0087739 | -0.0087739
| 12 | 2.1082 | 2.0685 | 2.0372 | 1.2156 | 2.1415 | 2.2661 | 1.6768 | 2.2214 | -0.0087739 | 0.75709 | 0.19866 | 0.33001 | -0.0087739 | -0.0087739 | -0.0087739
| 18 | 2.048 | 2.0571 | 2.1151 |  | 2.1878 | 2.1845 | 1.7289 |  | -0.0087739 | -0.0087739 | 0.15987 |  | -0.0087739 | -0.0087739 | -0.0087739
| 18 | 2.0532 | 2.0408 | 2.1101 |  | 2.2066 | 2.1823 | 1.7311 |  | -0.0087739 | -0.0087739 | 0.16613 |  | -0.0087739 | -0.0087739 | -0.0087739
| 18 | 2.0553 | 2.0413 | 2.1071 |  | 2.1778 | 2.1884 | -0.0087739 |  | -0.0087739 | -0.0087739 | 0.1232 |  | -0.0087739 | -0.0087739 | -0.0087739
| 24 | 2.0391 | 2.0269 | 2.2148 | 2.0265 | 2.1045 | 2.1911 | 1.7313 | 2.7841 | -0.0087739 | -0.0087739 | -0.0087739 | 1.3671 | -0.0087739 | -0.0087739
| 24 | 2.0508 | 2.0314 | 2.2102 | 2.0281 | 2.0937 | 2.1756 | 1.7305 | 2.7714 | -0.0087739 | -0.0087739 | -0.0087739 | 1.3626 | -0.0087739 | -0.0087739
| 24 | 2.0496 | 2.0279 | 2.2061 | 2.03 | 2.1001 | 2.1784 | -0.0087739 | 2.795 | -0.0087739 | -0.0087739 | -0.0087739 | 1.3664 | -0.0087739 | -0.0087739
| 48 | 0.55979 | 2.1109 |  | 1.8928 | 1.9959 | 2.0627 |  | 2.7969 | -0.0087739 |  |  | 0.98132 |  |  | -0.0087739
| 48 | 0.55715 | 2.1149 |  | 1.8952 | 2.0023 | 2.0558 |  | 2.7947 | -0.0087739 |  |  | 0.98628 |  |  | -0.0087739
| 48 | 0.54593 | 2.1132 |  | 1.8922 | 1.9912 | 2.0676 |  | 2.7914 | -0.0087739 |  |  | 0.97658 |  |  | -0.0087739
| Raw | CSB+Spaceru |  |  |  | CSBv+tSpacerv+t |  |  |  | Spacerv+t |  |  |  |  | VP
|  | G1B | G1C | G1E | G1F | G2B | G2C | G2D | G2E | G3A | G3B | G3C | G3D
time (min) | time (hr) | AB+Su.ii | AB+S.iii | AB+S.v | AB+S.vi | AB+SV2T2.ii | AB+SV2T2.iii | AB+SV2T2.iv | AB+SV2T2.v | SV2T2.i | SV2T2.ii | SV2T2.iii | SV2T2.iv | VP.i | VP.ii | VP.iii
5 | 0.083 | 0.98 | 209.329 | 438.982 | 1756.628 | 115.643 | 155.372 | 725.404 | 1480.395 | 9.895 | 36.074 | 50.965 | 64.516 | 7628.782 | 7802.841 | 0.98
5 | 0.083 | 0.98 | 206.651 | 628.151 | 1739.078 | 114.653 | 99.547 | 727.889 | 1489.326 | 10.175 | 35.979 | 50.462 | 63.211 | 7811.975 | 7461.587 | 0.98
5 | 0.083 | 0.98 | 197.532 | 622.532 | 1532.851 | 114.283 | 100.024 | 690.439 | 1435.4 | 9.554 | 35.946 | 51.48 | 75.198 | 7515.944 | 7456.986 | 0.98
10 | 0.167 | 17.778 | 118.83 | 355.403 | 1161.33 | 104.803 | 83.678 | 689.553 | 1545.515 | 7.522 | 38.117 | 40.937 | 46.293 | 5198.52 | 4453.828 | 0.98
10 | 0.167 | 17.598 | 119.529 | 355.784 | 1167.467 | 105.127 | 84.057 | 690.33 | 1575.516 | 7.09 | 37.304 | 40.314 | 44.471 | 5147.306 | 4455.158 | 0.98
10 | 0.167 | 18.017 | 121.162 | 370.965 | 1202.393 | 103.913 | 0.98 | 703.562 | 1582.915 | 7.292 | 37.932 | 41.174 | 46.581 | 5197.067 | 4442.14 | 0.98
15 | 0.25 | 14.896 | 26.641 | 480.451 | 857.017 | 94.557 | 73.845 | 667.63 | 1731.224 | 7.488 | 28.227 | 57.552 | 42.081 | 3333.629 | 2831.324 | 0.98
| 0.25 | 14.143 | 26.768 | 508.431 | 854.026 | 89.615 | 72.872 | 642.651 | 1768.111 | 7.296 | 28.16 | 58.684 | 46.853 | 3365.401 | 2881.125 | 0.98
| 0.25 | 13.938 | 26.765 | 494.034 | 856.695 | 89.849 | 71.861 | 620.444 | 1789.539 | 7.621 | 28.816 | 56.279 | 46.897 | 3408.065 | 2792.958 | 0.98
20 | 0.333 | 14.313 | 19.07 | 257.057 | 589.118 | 86.257 | 67.636 | 540.326 | 1695.594 | 5.781 | 21.606 | 30.257 | 31.847 | 2374.526 | 1811.193 | 236.451
| 0.333 | 14.19 | 18.937 | 251.926 | 609.306 | 86.714 | 66.282 | 542.219 | 1715.134 | 5.764 | 20.927 | 30.749 | 33.601 | 2382.916 | 1774.804 | 245.185
| 0.333 | 14.269 | 18.647 | 249.289 | 605.04 | 84.841 | 0.98 | 533.691 | 1729.01 | 5.913 | 21.653 | 31.123 | 31.696 | 2253.312 | 1800.969 | 233.782
25 | 0.417 | 12.54 | 15.089 | 161.325 | 418.879 | 76.758 | 61.203 | 467.787 | 1219.536 | 5.081 | 17.444 | 36.823 | 27.872 | 1335.373 | 923.957 | 38.14
| 0.417 | 12.662 | 14.748 | 153.992 | 431.211 | 76.543 | 0.98 | 457.154 | 1208.746 | 5.042 | 17.137 | 35.546 | 27.418 | 1348.087 | 942.098 | 36.987
| 0.417 | 12.515 | 14.92 | 165.312 | 447.715 | 77.652 | 0.98 | 464.22 | 1201.207 | 4.801 | 17.145 | 35.903 | 27.966 | 1345.818 | 916.849 | 37.51
30 | 0.5 | 12.71 | 15.174 | 77.462 | 341.836 | 87.674 | 0.98 | 355.751 | 898.652 | 5.444 | 12.34 | 14.801 | 26.524 | 753.796 | 402.567 | 5.098
| 0.5 | 12.796 | 15.293 | 46.137 | 340.097 | 84.632 | 0.98 | 363.456 | 884.988 | 5.461 | 12.634 | 15.007 | 26.134 | 748.017 | 391.356 | 4.503
| 0.5 | 12.644 | 17.183 | 31.464 | 317.483 | 85.65 | 0.98 | 358.447 | 875.605 | 5.223 | 12.192 | 15.269 | 26.322 | 755.255 | 390.593 | 5.173
35 | 0.583 | 13.459 | 16.917 | 0.98 | 244.985 | 81.305 | 0.98 | 227.139 | 613.209 | 3.007 | 15.495 | 17.441 | 15.034 | 416.569 | 19.222 | 0.98
| 0.583 | 13.654 | 16.832 | 70.012 | 250.054 | 80.11 | 0.98 | 225.881 | 615.287 | 2.841 | 15.777 | 17.735 | 15.199 | 403.321 | 20.085 | 0.98
| 0.583 | 13.553 | 17.154 | 33.762 | 248.068 | 80.83 | 0.98 | 224.002 | 559.803 | 2.87 | 15.073 | 17.133 | 15.545 | 407.994 | 17.188 | 0.98
40 | 0.667 | 13.654 | 16.656 | 8.251 | 232.628 | 80.915 | 0.98 | 143.34 | 356.569 | 1.086 | 9.233 | 23.592 | 13.699 | 193.282 | 138.165
| 0.667 | 13.623 | 16.811 | 8.112 | 239.338 | 81.549 | 0.98 | 147.324 | 351.478 | 1.166 | 8.506 | 23.45 | 13.701 | 198.144 | 134.804
| 0.667 | 13.48 | 16.585 | 8.031 | 238.593 | 82.075 | 0.98 | 146.565 | 344.737 | 1.324 | 8.878 | 24.135 | 13.762 | 195.793 | 136.576
45 | 0.75 | 13.777 | 15.443 | 8.123 | 156.556 | 76.56 | 0.98 | 61.104 | 213.454 | 2.818 | 9.591 | 13.176 | 10.272 | 142.245 | 117.845 | 0.98
| 0.75 | 13.969 | 15.539 | 8.075 | 157.751 | 76.659 | 0.98 | 55.891 | 216.869 | 2.845 | 9.535 | 13.463 | 10.289 | 142.971 | 117.468 | 0.98
| 0.75 | 13.753 | 15.39 | 8.063 | 158.238 | 79.535 | 0.98 | 62.781 | 224.48 | 2.609 | 9.419 | 13.414 | 10.245 | 145.124 | 121.728 | 0.98
50 | 0.833 | 13.39 | 15.446 | 7.96 | 128.955 | 74.406 | 0.98 | 22.681 | 124.13 | 0.98 | 10.253 | 15.982 | 9.934 | 63.303 | 82.379 | 0.98
| 0.833 | 13.455 | 15.446 | 7.933 | 130.09 | 74.732 | 0.98 | 17.07 | 126.156 | 0.98 | 9.706 | 15.615 | 10.008 | 67.764 | 81.942 | 0.98
| 0.833 | 13.447 | 15.427 | 7.88 | 130.901 | 72.823 | 0.98 | 19.891 | 127.094 | 0.98 | 10.187 | 15.857 | 10.007 | 65.107 | 84.776 | 0.98
55 | 0.917 | 12.966 | 15.951 | 50.451 | 119.741 | 83.206 | 0.98 | 0.98 | 69.051 | 0.98 | 12.861 | 10.886 | 5.89 | 0.98 | 54.434 | 0.98
| 0.917 | 13.582 | 16.327 | 50.704 | 121.983 | 81.169 | 0.98 | 103.857 | 68.014 | 0.98 | 12.591 | 10.866 | 5.663 | 0.98 | 56.598 | 0.98
| 0.917 | 13.486 | 15.333 | 49.913 | 123.091 | 82.522 | 0.98 | 104.468 | 68.723 | 0.98 | 13.034 | 10.841 | 5.738 | 0.98 | 54.68 | 0.98
60 | 1 | 16.204 | 110.953 | 55.454 | 104.089 | 86.208 | 0.98 | 0.98 | 1.445 | 0.98 | 9.435 | 15.176 | 3.811 | 0.98 | 34.741 | 2.449
| 1 | 16.119 | 110.266 | 56.284 | 103.572 | 86.809 | 0.98 | 0.98 | 4.407 | 0.98 | 9.674 | 15.088 | 3.912 | 0.98 | 34.354 | 2.583
| 1 | 16.192 | 110.387 | 56.524 | 104.938 | 85.155 | 0.98 | 52.142 | 2.125 | 0.98 | 9.344 | 15.023 | 3.833 | 0.98 | 34.64 | 2.553
| 2 | 15.415 | 19.184 | 60.365 | 11.343 | 85.607 | 0.98 | 37.26 | 35.395 | 0.98 | 0.98 | 3.523 | 1.629 | 0.98 | 0.98 | 0.98
| 2 | 15.748 | 19.37 | 60.288 | 11.589 | 87.056 | 0.98 | 0.98 | 35.078 | 0.98 | 0.98 | 3.55 | 1.57 | 0.98 | 0.98 | 0.98
| 2 | 15.646 | 19.496 | 59.851 | 11.541 | 86.529 | 0.98 | 0.98 | 35.374 | 0.98 | 0.98 | 3.718 | 1.503 | 0.98 | 0.98 | 0.98
| 3 | 27.117 | 27.793 | 51.311 | 15.761 | 117.84 | 141.686 | 0.98 | 77.439 | 0.98 | 0.98 | 1.463 | 3.002 | 0.98 | 0.98 | 0.98
| 3 | 27.345 | 27.93 | 52.806 | 15.602 | 116.229 | 165.5 | 0.98 | 77.018 | 0.98 | 0.98 | 1.327 | 2.85 | 0.98 | 0.98 | 0.98
| 3 | 26.786 | 27.648 | 52.101 | 15.646 | 118.613 | 138.526 | 0.98 | 80.402 | 0.98 | 0.98 | 1.167 | 2.866 | 0.98 | 0.98 | 0.98
| 4 | 129.16 | 29.908 | 48.168 | 17.888 | 119.591 | 130.566 | 0.98 | 93.005 | 0.98 | 0.98 | 5.301 | 3.638 | 0.98 | 0.98 | 0.98
| 4 | 129.646 | 29.666 | 49.107 | 17.652 | 117.355 | 129.778 | 0.98 | 92.822 | 0.98 | 0.98 | 5.457 | 3.477 | 0.98 | 0.98 | 0.98
| 4 | 129.924 | 29.857 | 49.402 | 17.746 | 118.826 | 153.533 | 60.919 | 91.604 | 0.98 | 0.98 | 5.517 | 3.635 | 0.98 | 0.98 | 0.98
| 5 | 0.98 | 155.28 | 27.717 | 18.62 | 163.155 | 288.118 | 76.268 | 99.483 | 0.98 | 0.98 | 3.848 | 3.48 | 0.98 | 0.98 | 0.98
| 5 | 173.022 | 161.584 | 27.452 | 18.468 | 164.343 | 278.884 | 0.98 | 98.466 | 0.98 | 0.98 | 3.819 | 3.325 | 0.98 | 0.98 | 0.98
| 5 | 171.8 | 156.547 | 27.287 | 18.424 | 162.286 | 261.305 | 75.588 | 103.825 | 0.98 | 0.98 | 3.871 | 3.301 | 0.98 | 0.98 | 0.98
| 6 | 171.689 | 158.709 | 41.683 | 11.74 | 155.847 | 251.336 | 83.39 | 71.208 | 0.98 | 0.98 | 5.939 | 0.98 | 0.98 | 0.98 | 0.98
| 6 | 172.26 | 161.06 | 40.701 | 11.618 | 146.765 | 248.269 | 84.071 | 70.642 | 0.98 | 0.98 | 6.205 | 0.98 | 0.98 | 0.98 | 0.98
| 6 | 170.706 | 160.263 | 41.03 | 11.604 | 150.085 | 243.231 | 0.98 | 0.98 | 0.98 | 0.98 | 6.053 | 0.98 | 0.98 | 0.98 | 0.98
| 7 | 50.236 | 135.521 | 78.943 | 13.246 | 154.043 | 238.436 | 91.463 | 78.346 | 0.98 | 0.98 | 4.977 | 0.98 | 0.98 | 0.98
| 7 | 50.352 | 135.96 | 79.787 | 13.27 | 148.648 | 239.831 | 89.507 | 77.119 | 0.98 | 0.98 | 5.277 | 0.98 | 0.98 | 0.98
| 7 | 49.451 | 138.083 | 79.603 | 13.287 | 147.36 | 228.973 | 87.995 | 77.384 | 0.98 | 0.98 | 5.228 | 0.98 | 0.98 | 0.98
| 8 | 139.733 | 127.59 | 97.176 |  | 142.532 | 189.709 | 82.761 |  | 0.98 | 19.133 | 3.88 |  | 0.98 | 0.98 | 0.98
| 8 | 142.113 | 131.109 | 96.321 |  | 142.736 | 188.702 | 83.09 |  | 0.98 | 19.04 | 3.83 |  | 0.98 | 0.98 | 0.98
| 8 | 142.868 | 130.265 | 94.797 |  | 138.148 | 181.876 | 0.98 |  | 0.98 | 19.069 | 3.82 |  | 0.98 | 0.98 | 0.98
| 12 | 126.843 | 117.358 | 108.479 | 16.545 | 140.925 | 184.731 | 0.98 | 133.092 | 0.98 | 6.016 | 1.468 | 2.118 | 0.98 | 0.98 | 0.98
| 12 | 128.016 | 116.515 | 107.388 | 16.477 | 142.164 | 182.482 | 47.501 | 130.5 | 0.98 | 5.993 | 1.421 | 2.066 | 0.98 | 0.98 | 0.98
| 12 | 128.293 | 117.093 | 108.946 | 16.429 | 138.505 | 184.56 | 47.509 | 166.502 | 0.98 | 5.716 | 1.58 | 2.138 | 0.98 | 0.98 | 0.98
| 18 | 111.695 | 114.063 | 130.355 |  | 154.109 | 152.938 | 53.573 |  | 0.98 | 0.98 | 1.445 |  | 0.98 | 0.98 | 0.98
| 18 | 113.032 | 109.846 | 128.843 |  | 160.922 | 152.172 | 53.836 |  | 0.98 | 0.98 | 1.466 |  | 0.98 | 0.98 | 0.98
| 18 | 113.581 | 109.973 | 127.96 |  | 150.583 | 154.327 | 0.98 |  | 0.98 | 0.98 | 1.328 |  | 0.98 | 0.98 | 0.98
| 24 | 109.41 | 106.38 | 163.969 | 106.299 | 127.199 | 155.272 | 53.86 | 608.289 | 0.98 | 0.98 | 0.98 | 23.285 | 0.98 | 0.98
| 24 | 112.414 | 107.507 | 162.259 | 106.687 | 124.083 | 149.844 | 53.765 | 590.735 | 0.98 | 0.98 | 0.98 | 23.045 | 0.98 | 0.98
| 24 | 112.093 | 106.626 | 160.746 | 107.161 | 125.926 | 150.785 | 0.98 | 623.667 | 0.98 | 0.98 | 0.98 | 23.249 | 0.98 | 0.98
| 48 | 3.629 | 129.085 |  | 78.126 | 99.064 | 115.52 |  | 626.518 | 0.98 |  |  | 9.579 |  |  | 0.98
| 48 | 3.607 | 130.298 |  | 78.553 | 100.53 | 113.717 |  | 623.257 | 0.98 |  |  | 9.689 |  |  | 0.98
| 48 | 3.515 | 129.775 |  | 78.022 | 97.985 | 116.837 |  | 618.545 | 0.98 |  |  | 9.475 |  |  | 0.98

## Tobramycin
TOB-PAO1 |  | CSBv+tSpaceru data was not used because it was clinically irrelevant | [44 empty] | Standard Error
| Log-scale | CSB+Spaceru |  |  |  | CSBv+tSpacerv+t |  |  |  | Spacerv+t |  |  |  |  | Geometric mean | CSB+Spaceru |  |  |  | CSBv+tSpacerv+t |  |  |  | Spacerv+t |  |  |  | Time (Days) | Time (Hr) | Average |  |  |  | Average |  |  |  |  | Average | [5 empty] | Flow Rate
|  | G1B | G1C | G1E | G1F | G2B | G2C | G2D | G2E | G3A | G3B | G3C | G3D |  |  | G1B | G1C | G1E | G1F | G2B | G2C | G2D | G2E | G3A | G3B | G3C | G3D |  |  | CSBv+tplusSpaceru | SD | SE | n | CSBv+tplusSpacerv+t | SD | SE | CI | n | Spacerv+t | SD | SE | CI | n | Min. Detection | Hour | mL/min
time (min) | time (hr) |  |  |  |  | CSB+S.ii | CSB+S.iii | CSB+S.iv | CSB+S.v | S.i | S.ii | S.iii | S.iv | time (min) | time (hr) |  |  |  |  | CSB+S.ii | CSB+S.iii | CSB+S.iv | CSB+S.v | S.i | S.ii | S.iii | S.iv | 0.0026042 | 0.0625 | 2.3011 | 0.62678 | 0.31339 | 4 | 2.4253 | 0.67599 | 0.338 | 0.65909 | 4 | 1.818 | 0.21018 | 0.10509 | 0.20597 | 4 | -0.0087739 | 0.0625 | 3.5
5 | 0.083 | 1.7266 | 2.0939 | 2.2339 | 3.1917 | 1.9426 | 1.8303 | 2.6652 | 3.2675 | 1.589 | 2.0923 | 1.8379 | 1.7556 | 5 | 0.083 | 1.7285 | 2.0941 | 2.1889 | 3.1929 | 1.9326 | 1.8283 | 2.6716 | 3.2685 | 1.5892 | 2.0935 | 1.8347 | 1.7548 | 0.0069583 | 0.167 | 2.2392 | 0.77377 | 0.38688 | 4 | 2.1821 | 0.80882 | 0.40441 | 0.7886 | 4 | 1.7483 | 0.25541 | 0.1277 | 0.2503 | 4 | -0.0087739 | 1 | 3.5
5 | 0.083 | 1.7295 | 2.0935 | 2.1776 | 3.1953 | 1.9275 | 1.827 | 2.6635 | 3.2745 | 1.5866 | 2.0931 | 1.8344 | 1.7559 | 10" | 0.167 | 1.2503 | 2.0786 | 2.5571 | 3.0707 | 1.5839 | 1.4284 | 2.6104 | 3.1058 | 1.4949 | 2.0735 | 1.8199 | 1.6048 | 0.010417 | 0.25 | 2.0523 | 0.89085 | 0.44542 | 4 | 2.0686 | 0.81804 | 0.40902 | 0.79759 | 4 | 1.6672 | 0.35498 | 0.17749 | 0.34788 | 4 | -0.0087739 | 1 | 3.13
5 | 0.083 | 1.7295 | 2.095 | 2.1551 | 3.1916 | 1.9277 | 1.8277 | 2.6862 | 3.2637 | 1.5919 | 2.095 | 1.8319 | 1.7528 | 15" | 0.25 | 1.1559 | 1.4269 | 2.6939 | 2.9324 | 1.4688 | 1.2843 | 2.5672 | 2.954 | 1.2508 | 2.0623 | 1.8332 | 1.5225 | 0.013875 | 0.333 | 1.9029 | 0.81054 | 0.40527 | 4 | 1.9715 | 0.82069 | 0.41035 | 0.80017 | 4 | 1.6023 | 0.37288 | 0.18644 | 0.36542 | 4 | -0.0087739 | 1.9999 | 3.13
10 | 0.167 | 1.2499 | 2.0749 | 2.5507 | 3.065 | 1.5863 | 1.4244 | 2.5999 | 3.1034 | 1.4987 | 2.0726 | 1.8201 | 1.6125 | 20" | 0.333 | 1.154 | 1.2761 | 2.4027 | 2.7789 | 1.361 | 1.1839 | 2.5124 | 2.8288 | 1.2371 | 2.0498 | 1.7613 | 1.3609 | 0.017375 | 0.417 | 1.7784 | 0.76236 | 0.38118 | 4 | 1.8619 | 0.81768 | 0.40884 | 0.79724 | 4 | 1.5049 | 0.43065 | 0.21532 | 0.42204 | 4 | -0.0087739 | 2 | 1.54
10 | 0.167 | 1.2455 | 2.0775 | 2.5512 | 3.0672 | 1.5807 | 1.4322 | 2.604 | 3.1047 | 1.5016 | 2.0733 | 1.8176 | 1.5998 | 25" | 0.417 | 1.0994 | 1.1737 | 2.2045 | 2.6359 | 1.2491 | 1.0715 | 2.4606 | 2.6663 | 1.0808 | 2.0341 | 1.6655 | 1.239 | 0.020833 | 0.5 | 1.6277 | 0.64812 | 0.32406 | 4 | 1.7903 | 0.77715 | 0.38857 | 0.75772 | 4 | 1.1995 | 0.17367 | 0.086837 | 0.1702 | 4 | -0.0087739 | 3 | 1.54
10 | 0.167 | 1.2557 | 2.0834 | 2.5693 | 3.08 | 1.5846 | 1.4286 | 2.6272 | 3.1093 | 1.4844 | 2.0747 | 1.8222 | 1.6023 | 30" | 0.5 | 1.1044 | 1.2002 | 1.6837 | 2.5224 | 1.208 | 1.0403 | 2.3496 | 2.5632 | 1.0639 | 1.0992 | 1.4483 | 1.1865 | 0.024292 | 0.583 | 1.47 | 0.61777 | 0.30889 | 4 | 1.7323 | 0.69881 | 0.3494 | 0.68134 | 4 | 1.1488 | 0.25128 | 0.12564 | 0.24625 | 4 | -0.0087739 | 3 | 1.02
| 0.25 | 1.1731 | 1.4256 | 2.6816 | 2.933 | 1.4653 | 1.288 | 2.5699 | 2.9496 | 1.2537 | 2.0607 | 1.8395 | 1.5244 | 35" | 0.583 | 1.1321 | 1.2296 | 1.1245 | 2.3939 | 1.1927 | 1.0726 | 2.2334 | 2.4305 | 0.93595 | 1.0842 | 1.5127 | 1.0622 | 0.027792 | 0.667 | 1.41 | 0.65624 | 0.32812 | 4 | 1.6725 | 0.64829 | 0.32414 | 0.63208 | 4 | 1.0638 | 0.27948 | 0.13974 | 0.27389 | 4 | -0.0087739 | 4 | 1.02
| 0.25 | 1.1505 | 1.4276 | 2.7062 | 2.9315 | 1.4745 | 1.283 | 2.5669 | 2.9517 | 1.252 | 2.0617 | 1.8343 | 1.5231 | 40" | 0.667 | 1.1331 | 1.2223 | 0.91013 | 2.3744 | 1.1728 | 1.0628 | 2.1179 | 2.3366 | 0.82029 | 1.0182 | 1.4645 | 0.95238 | 0.03125 | 0.75 | 1.3588 | 0.57236 | 0.28618 | 4 | 1.5975 | 0.58908 | 0.29454 | 0.57435 | 4 | 1.0156 | 0.25473 | 0.12736 | 0.24963 | 4 | -0.0087739 | 4 | 0.76
| 0.25 | 1.1442 | 1.4276 | 2.6938 | 2.9328 | 1.4667 | 1.2817 | 2.5649 | 2.9609 | 1.2466 | 2.0646 | 1.826 | 1.5201 | 45" | 0.75 | 1.1409 | 1.1891 | 0.90779 | 2.1973 | 1.1513 | 1.036 | 2.004 | 2.1987 | 0.80532 | 1.0006 | 1.3781 | 0.87836 | 0.034708 | 0.833 | 1.3324 | 0.53572 | 0.26786 | 4 | 1.5342 | 0.57137 | 0.28569 | 0.55709 | 4 | 0.93413 | 0.30566 | 0.15283 | 0.29954 | 4 | -0.0087739 | 5 | 0.76
| 0.333 | 1.1557 | 1.2804 | 2.41 | 2.7702 | 1.3601 | 1.1886 | 2.507 | 2.8288 | 1.24 | 2.0496 | 1.7573 | 1.3677 | 50" | 0.833 | 1.1281 | 1.1886 | 0.89896 | 2.1139 | 1.1246 | 0.97926 | 1.8774 | 2.1554 | 0.62725 | 1.0164 | 1.324 | 0.76886 | 0.038208 | 0.917 | 1.5282 | 0.45092 | 0.22546 | 4 | 1.4291 | 0.46111 | 0.23055 | 0.44958 | 4 | 0.90209 | 0.30259 | 0.1513 | 0.29654 | 4 | -0.0087739 | 5 | 0.61
| 0.333 | 1.152 | 1.2773 | 2.4013 | 2.7848 | 1.3588 | 1.1812 | 2.5114 | 2.8274 | 1.2397 | 2.0497 | 1.7682 | 1.355 | 55" | 0.917 | 1.1252 | 1.2004 | 1.702 | 2.0849 | 1.1455 | 1.069 | 1.419 | 2.0831 | 0.6067 | 1.0712 | 1.2399 | 0.69058 | 0.041667 | 1 | 1.7547 | 0.38762 | 0.19381 | 4 | 1.3869 | 0.43142 | 0.21571 | 0.42063 | 4 | 0.89584 | 0.3855 | 0.19275 | 0.37779 | 4 | -0.0087739 | 6 | 0.61
| 0.333 | 1.1544 | 1.2706 | 2.3967 | 2.7818 | 1.364 | 1.182 | 2.5188 | 2.8303 | 1.2317 | 2.0501 | 1.7582 | 1.36 | 60" | 1 | 1.2088 | 2.0435 | 1.7489 | 2.0179 | 1.146 | 1.074 | 1.3111 | 2.0167 | 0.5868 | 0.97073 | 1.4112 | 0.61461 | 0.083333 | 2 | 1.3299 | 0.31371 | 0.15685 | 4 | 1.0269 | 0.14187 | 0.070933 | 0.13832 | 4 | 0.63344 | 0.21596 | 0.10798 | 0.21164 | 4 | -0.0087739 | 6 | 0.5
| 0.417 | 1.0983 | 1.1787 | 2.2077 | 2.6221 | 1.2502 | 1.0651 | 2.4612 | 2.6608 | 1.0865 | 2.0333 | 1.6634 | 1.2459 | 2' | 2 | 1.1932 | 1.2867 | 1.7794 | 1.0603 | 1.1805 | 1.1093 | 0.9416 | 0.87615 | 0.4786 | 0.57594 | 0.95184 | 0.52737 | 0.125 | 3 | 1.4471 | 0.2132 | 0.1066 | 4 | 1.1702 | 0.30071 | 0.15035 | 0.29319 | 4 | 0.56002 | 0.24674 | 0.12337 | 0.2418 | 4 | -0.0087739 | 7 | 0.5
| 0.417 | 1.1025 | 1.1687 | 2.1875 | 2.6347 | 1.2476 | 1.0715 | 2.4588 | 2.6693 | 1.076 | 2.0348 | 1.671 | 1.2357 | 3' | 3 | 1.4327 | 1.4439 | 1.7166 | 1.1951 | 1.3774 | 1.3867 | 0.74425 | 1.1725 | 0.26182 | 0.51489 | 0.85963 | 0.60374 | 0.16667 | 4 | 1.6314 | 0.36758 | 0.18379 | 4 | 1.2454 | 0.15747 | 0.078736 | 0.15353 | 4 | 0.63257 | 0.33007 | 0.16504 | 0.32347 | 4 | -0.0087739 | 7 | 0.43
| 0.417 | 1.0974 | 1.1738 | 2.2183 | 2.651 | 1.2495 | 1.0778 | 2.4618 | 2.6688 | 1.0799 | 2.0341 | 1.6621 | 1.2355 | 4' | 4 | 2.1125 | 1.4744 | 1.6892 | 1.2495 | 1.3642 | 1.3643 | 1.0319 | 1.2212 | 0.28733 | 0.51295 | 1.0715 | 0.65848 | 0.20833 | 5 | 1.5989 | 0.41075 | 0.20537 | 4 | 1.4677 | 0.36049 | 0.18025 | 0.35148 | 4 | 0.70096 | 0.26457 | 0.13229 | 0.25928 | 4 | -0.0087739 | 8 | 0.43
| 0.5 | 1.1041 | 1.1811 | 1.8891 | 2.5338 | 1.208 | 1.0425 | 2.3738 | 2.5501 | 1.0639 | 1.0961 | 1.4551 | 1.1907 | 5' | 5 | 1.491 | 2.1981 | 1.4391 | 1.2673 | 1.7562 | 1.7933 | 1.083 | 1.2383 | 0.61078 | 0.44425 | 1.0691 | 0.67975 | 0.25 | 6 | 1.7798 | 0.55467 | 0.27733 | 4 | 1.4394 | 0.40512 | 0.20256 | 0.39499 | 4 | 0.66745 | 0.32562 | 0.16281 | 0.31911 | 4 | -0.0087739 | 8 | 0.38
| 0.5 | 1.1071 | 1.1845 | 1.664 | 2.5316 | 1.2114 | 1.0406 | 2.3376 | 2.5877 | 1.0645 | 1.1009 | 1.4509 | 1.187 | 6' | 6 | 2.2344 | 2.2041 | 1.6142 | 1.0665 | 1.7659 | 1.8125 | 1.1193 | 1.0597 | 0.62357 | 0.3878 | 1.1339 | 0.52457 | 0.29167 | 7 | 1.7143 | 0.43273 | 0.21636 | 4 | 1.4375 | 0.39548 | 0.19774 | 0.38559 | 4 | 0.749 | 0.3436 | 0.1718 | 0.33673 | 4 | -0.0087739 | 12 | 0.38
| 0.5 | 1.1019 | 1.2351 | 1.4978 | 2.5017 | 1.2045 | 1.0377 | 2.3375 | 2.5518 | 1.0633 | 1.1007 | 1.4388 | 1.1818 | 7' | 7 | 1.6991 | 2.1352 | 1.9001 | 1.1228 | 1.7657 | 1.7939 | 1.1052 | 1.0853 | 1.0256 | 0.38993 | 1.0599 | 0.52066 | 0.33333 | 8 | 2.0821 | 0.088213 | 0.05093 | 3 | 1.4953 | 0.39436 | 0.22769 | 0.44399 | 3 | 0.80873 | 0.20381 | 0.11767 | 0.23063 | 3 | -0.0087739 | 12 | 0.25
| 0.583 | 1.129 | 1.2283 | 0 | 2.3891 | 1.1935 | 1.0731 | 2.2586 | 2.4247 | 0.93892 | 1.0918 | 1.5169 | 1.0671 | 8' | 8 | 2.151 | 2.1128 | 1.9827 |  | 1.7346 | 1.7111 | 1.0401 |  | 0.68134 | 0.70105 | 1.0438 |  | 0.5 | 12 | 1.8565 | 0.4273 | 0.21365 | 4 | 1.1939 | 0.72589 | 0.36294 | 0.70774 | 4 | 0.68135 | 0.15303 | 0.076516 | 0.14997 | 4 | -0.0087739 | 18 | 0.25
| 0.583 | 1.1353 | 1.2261 | 1.8452 | 2.398 | 1.1956 | 1.0711 | 2.1879 | 2.4316 | 0.93323 | 1.0773 | 1.5118 | 1.0619 | 12' | 12 | 2.1062 | 2.0681 | 2.0345 | 1.2171 | 1.7324 | 1.7037 | 0.17915 | 1.1602 | 0.84852 | 0.51746 | 0.7676 | 0.59184 | 0.75 | 18 | 2.0698 | 0.035602 | 0.020555 | 3 | 1.2702 | 0.99308 | 0.57335 | 1.118 | 3 | 0.84071 | 0.30803 | 0.17784 | 0.34857 | 3 | -0.0087739 | 18 | 0.16
| 0.583 | 1.132 | 1.2344 | 1.5284 | 2.3946 | 1.1889 | 1.0737 | 2.2538 | 2.4351 | 0.93571 | 1.0836 | 1.5094 | 1.0576 | 18' | 18 | 2.0522 | 2.0464 | 2.1108 |  | 2.113 | 1.5222 | 0.17541 |  | 1.1736 | 0.56573 | 0.78283 |  | 1 | 24 | 2.0785 | 0.088356 | 0.044178 | 4 | 1.212 | 0.68953 | 0.34476 | 0.67229 | 4 | 0.94217 | 0.36866 | 0.18433 | 0.36128 | 4 | -0.0087739 | 24 | 0.16
| 0.667 | 1.1353 | 1.2216 | 0.91651 | 2.3667 | 1.1747 | 1.0641 | 2.1806 | 2.3364 | 0.82569 | 1.0215 | 1.4653 | 0.95303 | 24' | 24 | 2.0465 | 2.0287 | 2.2104 | 2.0282 | 1.5824 | 1.5142 | 0.17868 | 1.5727 | 1.0894 | 0.54863 | 0.74842 | 1.3822 | 2 | 48 | 1.5202 | 0.8437 | 0.48711 | 3 | 1.2403 | 0.10395 | 0.060014 | 0.11703 | 3 | 0.77093 | 0.45686 | 0.32305 | 0.63317 | 2 | -0.0087739 | 24 | 0.12
| 0.667 | 1.1343 | 1.2256 | 0.90913 | 2.379 | 1.1816 | 1.0647 | 2.0833 | 2.3349 | 0.81743 | 1.0165 | 1.4631 | 0.94443 | 48' | 48 | 0.55429 | 2.113 |  | 1.8934 | 1.2042 | 1.1592 |  | 1.3575 | 0.44788 |  |  | 1.094 | [17 empty] | 48 | 0.12
| 0.667 | 1.1297 | 1.2197 | 0.90477 | 2.3777 | 1.162 | 1.0597 | 2.0897 | 2.3385 | 0.81776 | 1.0167 | 1.4651 | 0.95966 | [31 empty] | 48 | 0.016
| 0.75 | 1.1392 | 1.1887 | 0.90972 | 2.1947 | 1.1547 | 1.0349 | 1.9714 | 2.1973 | 0.81485 | 1.0055 | 1.3759 | 0.88195 | Mean | CSB+Spaceru |  |  |  | CSBv+tSpacerv+t |  |  |  | Spacerv+t | [21 empty] | 120 | 0.016
| 0.75 | 1.1452 | 1.1914 | 0.90714 | 2.198 | 1.1515 | 1.0355 | 2.0682 | 2.197 | 0.80557 | 1.0011 | 1.379 | 0.87616 |  | G1B | G1C | G1E | G1F | G2B | G2C | G2D | G2E | G3A | G3B | G3C | G3D |  |  |  | Log Dif | [14 empty] | 120 | 0.016
| 0.75 | 1.1384 | 1.1872 | 0.9065 | 2.1993 | 1.1477 | 1.0376 | 1.9724 | 2.2018 | 0.79553 | 0.99511 | 1.3793 | 0.87697 | Time (hour) |  |  |  |  | CSB+S.ii | CSB+S.iii | CSB+S.iv | CSB+S.v | S.i | S.ii | S.iii | S.iv |  |  |  | Time (Hr) | Average | Average | [12 empty] | 192 | 0.016
| 0.833 | 1.1268 | 1.1888 | 0.90091 | 2.1104 | 1.1228 | 0.97882 | 1.8702 | 2.1511 | 0.62767 | 1.0161 | 1.3197 | 0.76908 | 0.063 | 53.521 | 124.1977 | 154.9297 | 1559.123 | 85.6323 | 67.3487 | 469.6073 | 1855.9733 | 38.8303 | 124.012 | 68.349 | 56.8573 |  |  |  |  | CSBv+tplusSpacerv+t | Spacerv+t | Difference | [11 empty] | 264 | 0.016
| 0.833 | 1.1289 | 1.1888 | 0.89944 | 2.1142 | 1.1324 | 0.97964 | 1.8798 | 2.1571 | 0.63053 | 1.0164 | 1.325 | 0.76716 | 0.167 | 17.7977 | 119.8403 | 360.7173 | 1177.0633 | 38.359 | 26.8193 | 407.917 | 1275.8567 | 31.2577 | 118.4463 | 66.0613 | 40.2603 |  |  |  | 0.0625 | 2.4253 | 1.818 | 0.60723 | [11 empty] | 336 | 0.016
| 0.833 | 1.1286 | 1.1883 | 0.89653 | 2.1169 | 1.1185 | 0.97932 | 1.8823 | 2.1579 | 0.62356 | 1.0167 | 1.3274 | 0.77034 | 0.25 | 14.3257 | 26.7247 | 494.3053 | 855.9127 | 29.4323 | 19.2427 | 369.1743 | 899.641 | 17.8157 | 115.433 | 68.1213 | 33.307 |  |  |  | 0.167 | 2.1821 | 1.7483 | 0.43382 | [11 empty] | 408 | 0.016
| 0.917 | 1.1128 | 1.2028 | 1.7029 | 2.0782 | 1.147 | 1.069 | 1.4139 | 2.0826 | 0.60498 | 1.0718 | 1.2382 | 0.69073 | 0.333 | 14.2573 | 18.8847 | 252.7573 | 601.1547 | 22.9613 | 15.2737 | 325.4033 | 674.295 | 17.265 | 112.147 | 57.714 | 22.9583 |  |  |  | 0.25 | 2.0686 | 1.6672 | 0.40136 | [11 empty] | 480 | 0.016
| 0.917 | 1.133 | 1.2129 | 1.705 | 2.0863 | 1.1496 | 1.0679 | 1.4183 | 2.0827 | 0.61045 | 1.0725 | 1.2416 | 0.69002 | 0.417 | 12.5723 | 14.919 | 160.2097 | 432.6017 | 17.746 | 11.7903 | 288.8207 | 463.7973 | 12.0453 | 108.164 | 46.2923 | 17.3403 |  |  |  | 0.333 | 1.9715 | 1.6023 | 0.36927 | [11 empty] | 552 | 0.016
| 0.917 | 1.1299 | 1.1856 | 1.6982 | 2.0902 | 1.14 | 1.0701 | 1.4249 | 2.0839 | 0.60466 | 1.0694 | 1.2397 | 0.69099 | 0.5 | 12.7167 | 15.8833 | 51.6877 | 333.1387 | 16.1433 | 10.972 | 223.8643 | 366.0707 | 11.585 | 12.5677 | 28.0747 | 15.3647 |  |  |  | 0.417 | 1.8619 | 1.5049 | 0.35703 | [11 empty] | 624 | 0.016
| 1 | 1.2096 | 2.0451 | 1.7439 | 2.0174 | 1.1474 | 1.0751 | 1.3185 | 2.0174 | 0.58365 | 0.96478 | 1.4131 | 0.61658 | 0.583 | 13.5553 | 16.9677 | 34.918 | 247.7023 | 15.5833 | 11.8203 | 171.638 | 269.4703 | 8.629 | 12.1413 | 32.5643 | 11.54 |  |  |  | 0.5 | 1.7903 | 1.1995 | 0.5908 | [11 empty] | 696 | 0.016
| 1 | 1.2073 | 2.0424 | 1.7504 | 2.0152 | 1.1506 | 1.0775 | 1.3003 | 2.0151 | 0.59162 | 0.974 | 1.4125 | 0.61469 | 0.667 | 13.5857 | 16.684 | 8.1313 | 236.853 | 14.8887 | 11.5563 | 131.8823 | 217.078 | 6.6117 | 10.429 | 29.1387 | 8.9623 |  |  |  | 0.583 | 1.7323 | 1.1488 | 0.58353 | [11 empty] | 768 | 0.016
| 1 | 1.2093 | 2.0429 | 1.7522 | 2.0209 | 1.14 | 1.0693 | 1.3145 | 2.0176 | 0.58512 | 0.9734 | 1.4081 | 0.61257 | 0.75 | 13.833 | 15.4573 | 8.087 | 157.515 | 14.1677 | 10.8647 | 101.4877 | 158.019 | 6.3883 | 10.0137 | 23.8817 | 7.5573 |  |  |  | 0.667 | 1.6725 | 1.0638 | 0.60867 | [11 empty] | 840 | 0.016
| 2 | 1.1879 | 1.2829 | 1.7808 | 1.0547 | 1.183 | 1.1107 | 0.93852 | 0.87355 | 0.46835 | 0.5841 | 0.95607 | 0.52686 | 0.833 | 13.4307 | 15.4397 | 7.9243 | 129.982 | 13.3237 | 9.5337 | 75.4147 | 143.0133 | 4.239 | 10.3847 | 21.0883 | 5.873 |  |  |  | 0.75 | 1.5975 | 1.0156 | 0.58192 | [11 empty] | 912 | 0.016
| 2 | 1.1972 | 1.2871 | 1.7802 | 1.064 | 1.1795 | 1.1093 | 0.93947 | 0.87558 | 0.48629 | 0.57692 | 0.95516 | 0.52673 | 0.917 | 13.3447 | 15.8703 | 50.356 | 121.605 | 13.981 | 11.7217 | 26.245 | 121.0763 | 4.043 | 11.7817 | 17.3727 | 4.9043 |  |  |  | 0.833 | 1.5342 | 0.93413 | 0.60003 | [11 empty] | 984 | 0.016
| 2 | 1.1944 | 1.2899 | 1.7771 | 1.0622 | 1.179 | 1.1078 | 0.9468 | 0.87933 | 0.48116 | 0.56679 | 0.94429 | 0.52853 | 1 | 16.1717 | 110.5353 | 56.0873 | 104.1997 | 13.9957 | 11.8577 | 20.4727 | 103.923 | 3.862 | 9.3487 | 25.7777 | 4.1173 |  |  |  | 0.917 | 1.4291 | 0.90209 | 0.52706
| 3 | 1.4332 | 1.4439 | 1.7102 | 1.1976 | 1.3762 | 1.3873 | 0.74107 | 1.1673 | 0.26245 | 0.51601 | 0.86279 | 0.60681 | 2 | 15.603 | 19.35 | 60.168 | 11.491 | 15.1533 | 12.8603 | 8.742 | 7.519 | 3.0107 | 3.767 | 8.951 | 3.368 |  |  |  | 1 | 1.3869 | 0.89584 | 0.4911
| 3 | 1.4369 | 1.4461 | 1.7227 | 1.1932 | 1.3784 | 1.3845 | 0.74186 | 1.1792 | 0.2615 | 0.51733 | 0.85818 | 0.5978 | 3 | 27.0827 | 27.7903 | 52.0727 | 15.6697 | 23.847 | 24.364 | 5.5497 | 14.8777 | 1.8273 | 3.2727 | 7.2383 | 4.0157 |  |  |  | 2 | 1.0269 | 0.63344 | 0.39344
| 3 | 1.4279 | 1.4417 | 1.7168 | 1.1944 | 1.3777 | 1.3884 | 0.74981 | 1.1711 | 0.2615 | 0.51135 | 0.85794 | 0.6066 | 4 | 129.5767 | 29.8103 | 48.8923 | 17.762 | 23.1323 | 23.1347 | 10.7633 | 16.6417 | 1.938 | 3.258 | 11.7907 | 4.555 |  |  |  | 3 | 1.1702 | 0.56002 | 0.61021
| 4 | 2.1111 | 1.4758 | 1.6828 | 1.2526 | 1.3694 | 1.3627 | 1.0357 | 1.2196 | 0.2938 | 0.51001 | 1.0713 | 0.6621 | 5 | 115.2673 | 157.8037 | 27.4853 | 18.504 | 57.0473 | 62.133 | 12.1073 | 17.3113 | 4.0813 | 2.7813 | 11.724 | 4.7837 |  |  |  | 4 | 1.2454 | 0.63257 | 0.61283
| 4 | 2.1128 | 1.4723 | 1.6911 | 1.2468 | 1.3617 | 1.3688 | 1.0295 | 1.2203 | 0.28443 | 0.51561 | 1.0769 | 0.65896 | 6 | 171.5517 | 160.0107 | 41.138 | 11.654 | 58.3347 | 64.9333 | 13.1643 | 11.4747 | 4.204 | 2.4423 | 13.612 | 3.3463 |  |  |  | 5 | 1.4677 | 0.70096 | 0.76676
| 4 | 2.1137 | 1.475 | 1.6937 | 1.2491 | 1.3615 | 1.3613 | 1.0306 | 1.2237 | 0.28375 | 0.51322 | 1.0664 | 0.65437 | 7 | 50.013 | 136.5213 | 79.4443 | 13.2677 | 58.3133 | 62.221 | 12.741 | 12.1693 | 10.6067 | 2.4543 | 11.4777 | 3.3163 |  |  |  | 6 | 1.4394 | 0.66745 | 0.77191
| 5 | 0 | 2.1911 | 1.4427 | 1.27 | 1.7608 | 1.7929 | 1.0868 | 1.243 | 0.61268 | 0.44638 | 1.0693 | 0.6766 | 8 | 141.5713 | 129.6547 | 96.098 |  | 54.31 | 51.4253 | 10.969 |  | 4.8013 | 5.0243 | 11.061 |  |  |  |  | 7 | 1.4375 | 0.749 | 0.68851
| 5 | 2.2381 | 2.2084 | 1.4386 | 1.2664 | 1.7553 | 1.796 | 1.0775 | 1.2386 | 0.60433 | 0.44545 | 1.0695 | 0.68431 | 12 | 127.7173 | 116.9887 | 108.271 | 16.4837 | 54.0063 | 50.544 | 1.5107 | 14.461 | 7.1203 | 3.292 | 5.856 | 3.907 |  |  |  | 8 | 1.4953 | 0.80873 | 0.68657
| 5 | 2.235 | 2.1946 | 1.436 | 1.2654 | 1.7526 | 1.7911 | 1.0848 | 1.2333 | 0.61532 | 0.44091 | 1.0684 | 0.67834 | 18 | 112.7693 | 111.294 | 129.0527 |  | 129.7193 | 33.283 | 1.4977 |  | 14.914 | 3.679 | 6.065 |  |  |  |  | 12 | 1.1939 | 0.68135 | 0.51251
| 6 | 2.2347 | 2.2006 | 1.62 | 1.0697 | 1.7716 | 1.8131 | 1.1294 | 1.0608 | 0.63609 | 0.3897 | 1.1402 | 0.52453 | 24 | 111.3057 | 106.8377 | 162.3247 | 106.7157 | 38.2297 | 32.675 | 1.509 | 37.3923 | 12.33 | 3.537 | 5.603 | 24.111 |  |  |  | 18 | 1.2702 | 0.84071 | 0.4295
| 6 | 2.2362 | 2.207 | 1.6096 | 1.0651 | 1.7623 | 1.8114 | 1.1151 | 1.0592 | 0.61679 | 0.38703 | 1.1368 | 0.52336 | 48 | 3.5837 | 129.7193 |  | 78.2337 | 16.0023 | 14.4287 |  | 22.7783 | 2.8047 |  |  | 12.4167 |  |  |  | 24 | 1.212 | 0.94217 | 0.26982
| 6 | 2.2322 | 2.2048 | 1.6131 | 1.0646 | 1.7638 | 1.8129 | 1.1135 | 1.0592 | 0.61784 | 0.38668 | 1.1247 | 0.52582 | [16 empty] | 48 | 1.2403 | 0.77093 | 0.46937
| 7 | 1.701 | 2.132 | 1.8973 | 1.1221 | 1.7685 | 1.8046 | 1.1067 | 1.0852 | 1.0305 | 0.38863 | 1.0608 | 0.52205
| 7 | 1.702 | 2.1334 | 1.9019 | 1.1229 | 1.7717 | 1.7928 | 1.1022 | 1.0863 | 1.0243 | 0.39023 | 1.0592 | 0.52009
| 7 | 1.6942 | 2.1401 | 1.9009 | 1.1234 | 1.757 | 1.7841 | 1.1067 | 1.0843 | 1.0219 | 0.39094 | 1.0596 | 0.51983
| 8 | 2.1453 | 2.1058 | 1.9876 |  | 1.7147 | 1.7167 | 1.0347 |  | 0.67752 | 0.69496 | 1.0458
| 8 | 2.1526 | 2.1176 | 1.9837 |  | 1.7503 | 1.7128 | 1.0481 |  | 0.6878 | 0.70165 | 1.0428
| 8 | 2.1549 | 2.1148 | 1.9768 |  | 1.7389 | 1.7039 | 1.0377 |  | 0.6787 | 0.70655 | 1.0428
| 12 | 2.1033 | 2.0695 | 2.0353 | 1.2187 | 1.7326 | 1.7046 | 0.18241 | 1.1578 | 0.80996 | 0.51574 | 0.76723 | 0.59528
| 12 | 2.1073 | 2.0664 | 2.031 | 1.2169 | 1.7341 | 1.7056 | 0.1807 | 1.1648 | 0.80509 | 0.51851 | 0.76664 | 0.59017
| 12 | 2.1082 | 2.0685 | 2.0372 | 1.2156 | 1.7306 | 1.7008 | 0.17435 | 1.158 | 0.93049 | 0.51812 | 0.76893 | 0.59006
| 18 | 2.048 | 2.0571 | 2.1151 |  | 2.1109 | 1.5237 | 0.17926 |  | 1.1778 | 0.56455 | 0.78305
| 18 | 2.0532 | 2.0408 | 2.1101 |  | 2.1149 | 1.523 | 0.17348 |  | 1.1734 | 0.56761 | 0.78383
| 18 | 2.0553 | 2.0413 | 2.1071 |  | 2.1132 | 1.5199 | 0.17348 |  | 1.1695 | 0.56502 | 0.78161
| 24 | 2.0391 | 2.0269 | 2.2148 | 2.0265 | 1.583 | 1.5146 | 0.18127 | 1.5816 | 1.1389 | 0.54753 | 0.74733 | 1.3793
| 24 | 2.0508 | 2.0314 | 2.2102 | 2.0281 | 1.5832 | 1.5134 | 0.1784 | 1.5776 | 1.0526 | 0.55035 | 0.7485 | 1.3867
| 24 | 2.0496 | 2.0279 | 2.2061 | 2.03 | 1.5811 | 1.5146 | 0.17638 | 1.5589 | 1.0767 | 0.54802 | 0.74943 | 1.3806
| 48 | 0.55979 | 2.1109 |  | 1.8928 | 1.2024 | 1.1585 |  | 1.3655 | 0.44855 |  |  | 1.1012
| 48 | 0.55715 | 2.1149 |  | 1.8952 | 1.2049 | 1.1616 |  | 1.3539 | 0.44607 |  |  | 1.0921
| 48 | 0.54593 | 2.1132 |  | 1.8922 | 1.2052 | 1.1575 |  | 1.353 | 0.44902 |  |  | 1.0886
| Raw Data | CSB+Spaceru |  |  |  | CSBv+tSpacerv+t |  |  |  | Spacerv+t
|  | G1B | G1C | G1E | G1F | G2B | G2C | G2D | G2E | G3A | G3B | G3C | G3D
time (min) | time (hr) | AB+Su.ii | AB+S.iii | AB+S.v | AB+S.vi | AB+SV2T2.ii | AB+SV2T2.iii | AB+SV2T2.iv | AB+SV2T2.v | SV2T2.i | SV2T2.ii | SV2T2.iii | SV2T2.iv
5 | 0.083 | 53.28 | 124.124 | 171.347 | 1554.985 | 87.611 | 67.655 | 462.554 | 1851.237 | 38.816 | 123.672 | 68.845 | 56.968
5 | 0.083 | 53.642 | 124.014 | 150.533 | 1567.813 | 84.617 | 67.146 | 460.77 | 1881.409 | 38.597 | 123.917 | 68.298 | 57.001
5 | 0.083 | 53.641 | 124.455 | 142.909 | 1554.571 | 84.669 | 67.245 | 485.498 | 1835.274 | 39.078 | 124.447 | 67.904 | 56.603
10 | 0.167 | 17.778 | 118.83 | 355.403 | 1161.33 | 38.578 | 26.573 | 398.038 | 1268.835 | 31.531 | 118.187 | 66.079 | 40.969
10 | 0.167 | 17.598 | 119.529 | 355.784 | 1167.467 | 38.079 | 27.054 | 401.829 | 1272.621 | 31.737 | 118.384 | 65.705 | 39.788
10 | 0.167 | 18.017 | 121.162 | 370.965 | 1202.393 | 38.42 | 26.831 | 423.884 | 1286.114 | 30.505 | 118.768 | 66.4 | 40.024
| 0.25 | 14.896 | 26.641 | 480.451 | 857.017 | 29.192 | 19.41 | 371.459 | 890.426 | 17.937 | 115.005 | 69.097 | 33.454
| 0.25 | 14.143 | 26.768 | 508.431 | 854.026 | 29.819 | 19.188 | 368.852 | 894.671 | 17.865 | 115.259 | 68.276 | 33.35
| 0.25 | 13.938 | 26.765 | 494.034 | 856.695 | 29.286 | 19.13 | 367.212 | 913.826 | 17.645 | 116.035 | 66.991 | 33.117
| 0.333 | 14.313 | 19.07 | 257.057 | 589.118 | 22.916 | 15.44 | 321.403 | 674.272 | 17.379 | 112.094 | 57.187 | 23.321
| 0.333 | 14.19 | 18.937 | 251.926 | 609.306 | 22.846 | 15.176 | 324.627 | 672.018 | 17.365 | 112.12 | 58.643 | 22.646
| 0.333 | 14.269 | 18.647 | 249.289 | 605.04 | 23.122 | 15.205 | 330.18 | 676.595 | 17.051 | 112.227 | 57.312 | 22.908
| 0.417 | 12.54 | 15.089 | 161.325 | 418.879 | 17.791 | 11.618 | 289.224 | 457.938 | 12.205 | 107.979 | 46.066 | 17.616
| 0.417 | 12.662 | 14.748 | 153.992 | 431.211 | 17.683 | 11.79 | 287.637 | 466.979 | 11.912 | 108.34 | 46.877 | 17.206
| 0.417 | 12.515 | 14.92 | 165.312 | 447.715 | 17.764 | 11.963 | 289.601 | 466.475 | 12.019 | 108.173 | 45.934 | 17.199
| 0.5 | 12.71 | 15.174 | 77.462 | 341.836 | 16.143 | 11.028 | 236.497 | 354.915 | 11.584 | 12.476 | 28.518 | 15.512
| 0.5 | 12.796 | 15.293 | 46.137 | 340.097 | 16.272 | 10.98 | 217.586 | 387.029 | 11.602 | 12.616 | 28.242 | 15.382
| 0.5 | 12.644 | 17.183 | 31.464 | 317.483 | 16.015 | 10.908 | 217.51 | 356.268 | 11.569 | 12.611 | 27.464 | 15.2
| 0.583 | 13.459 | 16.917 | 0.98 | 244.985 | 15.612 | 11.834 | 181.375 | 265.903 | 8.688 | 12.353 | 32.88 | 11.67
| 0.583 | 13.654 | 16.832 | 70.012 | 250.054 | 15.689 | 11.778 | 154.131 | 270.158 | 8.575 | 11.949 | 32.496 | 11.533
| 0.583 | 13.553 | 17.154 | 33.762 | 248.068 | 15.449 | 11.849 | 179.408 | 272.35 | 8.624 | 12.122 | 32.317 | 11.417
| 0.667 | 13.654 | 16.656 | 8.251 | 232.628 | 14.951 | 11.59 | 151.567 | 216.962 | 6.694 | 10.508 | 29.192 | 8.975
| 0.667 | 13.623 | 16.811 | 8.112 | 239.338 | 15.193 | 11.606 | 121.138 | 216.235 | 6.568 | 10.387 | 29.045 | 8.799
| 0.667 | 13.48 | 16.585 | 8.031 | 238.593 | 14.522 | 11.473 | 122.942 | 218.037 | 6.573 | 10.392 | 29.179 | 9.113
| 0.75 | 13.777 | 15.443 | 8.123 | 156.556 | 14.28 | 10.837 | 93.618 | 157.515 | 6.529 | 10.127 | 23.765 | 7.62
| 0.75 | 13.969 | 15.539 | 8.075 | 157.751 | 14.173 | 10.852 | 116.992 | 157.395 | 6.391 | 10.026 | 23.933 | 7.519
| 0.75 | 13.753 | 15.39 | 8.063 | 158.238 | 14.05 | 10.905 | 93.853 | 159.147 | 6.245 | 9.888 | 23.947 | 7.533
| 0.833 | 13.39 | 15.446 | 7.96 | 128.955 | 13.269 | 9.524 | 74.16 | 141.621 | 4.243 | 10.377 | 20.878 | 5.876
| 0.833 | 13.455 | 15.446 | 7.933 | 130.09 | 13.564 | 9.542 | 75.828 | 143.583 | 4.271 | 10.386 | 21.133 | 5.85
| 0.833 | 13.447 | 15.427 | 7.88 | 130.901 | 13.138 | 9.535 | 76.256 | 143.836 | 4.203 | 10.391 | 21.254 | 5.893
| 0.917 | 12.966 | 15.951 | 50.451 | 119.741 | 14.027 | 11.721 | 25.933 | 120.951 | 4.027 | 11.797 | 17.308 | 4.906
| 0.917 | 13.582 | 16.327 | 50.704 | 121.983 | 14.111 | 11.692 | 26.2 | 120.971 | 4.078 | 11.816 | 17.442 | 4.898
| 0.917 | 13.486 | 15.333 | 49.913 | 123.091 | 13.805 | 11.752 | 26.602 | 121.307 | 4.024 | 11.732 | 17.368 | 4.909
| 1 | 16.204 | 110.953 | 55.454 | 104.089 | 14.04 | 11.888 | 20.82 | 104.096 | 3.834 | 9.221 | 25.891 | 4.136
| 1 | 16.119 | 110.266 | 56.284 | 103.572 | 14.144 | 11.954 | 19.968 | 103.538 | 3.905 | 9.419 | 25.852 | 4.118
| 1 | 16.192 | 110.387 | 56.524 | 104.938 | 13.803 | 11.731 | 20.63 | 104.135 | 3.847 | 9.406 | 25.59 | 4.098
| 2 | 15.415 | 19.184 | 60.365 | 11.343 | 15.241 | 12.902 | 8.68 | 7.474 | 2.94 | 3.838 | 9.038 | 3.364
| 2 | 15.748 | 19.37 | 60.288 | 11.589 | 15.118 | 12.862 | 8.699 | 7.509 | 3.064 | 3.775 | 9.019 | 3.363
| 2 | 15.646 | 19.496 | 59.851 | 11.541 | 15.101 | 12.817 | 8.847 | 7.574 | 3.028 | 3.688 | 8.796 | 3.377
| 3 | 27.117 | 27.793 | 51.311 | 15.761 | 23.779 | 24.397 | 5.509 | 14.699 | 1.83 | 3.281 | 7.291 | 4.044
| 3 | 27.345 | 27.93 | 52.806 | 15.602 | 23.899 | 24.24 | 5.519 | 15.107 | 1.826 | 3.291 | 7.214 | 3.961
| 3 | 26.786 | 27.648 | 52.101 | 15.646 | 23.863 | 24.455 | 5.621 | 14.827 | 1.826 | 3.246 | 7.21 | 4.042
| 4 | 129.16 | 29.908 | 48.168 | 17.888 | 23.411 | 23.052 | 10.857 | 16.579 | 1.967 | 3.236 | 11.784 | 4.593
| 4 | 129.646 | 29.666 | 49.107 | 17.652 | 22.997 | 23.376 | 10.702 | 16.607 | 1.925 | 3.278 | 11.937 | 4.56
| 4 | 129.924 | 29.857 | 49.402 | 17.746 | 22.989 | 22.976 | 10.731 | 16.739 | 1.922 | 3.26 | 11.651 | 4.512
| 5 | 0.98 | 155.28 | 27.717 | 18.62 | 57.649 | 62.078 | 12.212 | 17.498 | 4.099 | 2.795 | 11.731 | 4.749
| 5 | 173.022 | 161.584 | 27.452 | 18.468 | 56.927 | 62.512 | 11.953 | 17.323 | 4.021 | 2.789 | 11.735 | 4.834
| 5 | 171.8 | 156.547 | 27.287 | 18.424 | 56.566 | 61.809 | 12.157 | 17.113 | 4.124 | 2.76 | 11.706 | 4.768
| 6 | 171.689 | 158.709 | 41.683 | 11.74 | 59.097 | 65.021 | 13.47 | 11.504 | 4.326 | 2.453 | 13.809 | 3.346
| 6 | 172.26 | 161.06 | 40.701 | 11.618 | 57.852 | 64.774 | 13.036 | 11.46 | 4.138 | 2.438 | 13.702 | 3.337
| 6 | 170.706 | 160.263 | 41.03 | 11.604 | 58.055 | 65.005 | 12.987 | 11.46 | 4.148 | 2.436 | 13.325 | 3.356
| 7 | 50.236 | 135.521 | 78.943 | 13.246 | 58.678 | 63.771 | 12.786 | 12.168 | 10.727 | 2.447 | 11.502 | 3.327
| 7 | 50.352 | 135.96 | 79.787 | 13.27 | 59.112 | 62.06 | 12.652 | 12.198 | 10.576 | 2.456 | 11.46 | 3.312
| 7 | 49.451 | 138.083 | 79.603 | 13.287 | 57.15 | 60.832 | 12.785 | 12.142 | 10.517 | 2.46 | 11.471 | 3.31
| 8 | 139.733 | 127.59 | 97.176 |  | 51.842 | 52.089 | 10.831 |  | 4.759 | 4.954 | 11.112
| 8 | 142.113 | 131.109 | 96.321 |  | 56.277 | 51.617 | 11.17 |  | 4.873 | 5.031 | 11.035
| 8 | 142.868 | 130.265 | 94.797 |  | 54.811 | 50.57 | 10.906 |  | 4.772 | 5.088 | 11.036
| 12 | 126.843 | 117.358 | 108.479 | 16.545 | 54.03 | 50.649 | 1.522 | 14.381 | 6.456 | 3.279 | 5.851 | 3.938
| 12 | 128.016 | 116.515 | 107.388 | 16.477 | 54.212 | 50.772 | 1.516 | 14.615 | 6.384 | 3.3 | 5.843 | 3.892
| 12 | 128.293 | 117.093 | 108.946 | 16.429 | 53.777 | 50.211 | 1.494 | 14.387 | 8.521 | 3.297 | 5.874 | 3.891
| 18 | 111.695 | 114.063 | 130.355 |  | 129.085 | 33.399 | 1.511 |  | 15.06 | 3.669 | 6.068
| 18 | 113.032 | 109.846 | 128.843 |  | 130.298 | 33.342 | 1.491 |  | 14.908 | 3.695 | 6.079
| 18 | 113.581 | 109.973 | 127.96 |  | 129.775 | 33.108 | 1.491 |  | 14.774 | 3.673 | 6.048
| 24 | 109.41 | 106.38 | 163.969 | 106.299 | 38.281 | 32.707 | 1.518 | 38.156 | 13.77 | 3.528 | 5.589 | 23.95
| 24 | 112.414 | 107.507 | 162.259 | 106.687 | 38.296 | 32.616 | 1.508 | 37.807 | 11.287 | 3.551 | 5.604 | 24.364
| 24 | 112.093 | 106.626 | 160.746 | 107.161 | 38.112 | 32.702 | 1.501 | 36.214 | 11.933 | 3.532 | 5.616 | 24.019
| 48 | 3.629 | 129.085 |  | 78.126 | 15.938 | 14.405 |  | 23.202 | 2.809 |  |  | 12.625
| 48 | 3.607 | 130.298 |  | 78.553 | 16.028 | 14.509 |  | 22.589 | 2.793 |  |  | 12.361
| 48 | 3.515 | 129.775 |  | 78.022 | 16.041 | 14.372 |  | 22.544 | 2.812 |  |  | 12.264

## AUC
TOB
|  | CSBv+tSpacerv+t - 2 | CSBv+tSpacerv+t - 12 | CSBv+tSpacerv+t - 24 | CSBv+tSpacerv+t - 48 | Spacerv+t - 2 | Spacerv+t - 12 | Spacerv+t - 24 | Spacerv+t - 48
| 1 | 35.14 | 507.2 | 1562 | 2213 | 15.22 | 65.62 | 213.4 | 395
| 2 | 27.32 | 500.2 | 949.5 | 1515 | 55.63 | 89.95 | 132.5 | 132.5
| 3 | 212 | 301.1 | 319.2 | 319.2 | 53.83 | 153.5 | 224.3 | 224.3
| 4 | 520.4 | 657.1 | 968.2 | 1690 | 20.31 | 58.44 | 226.6 | 664.9
|  | 2 hours | 12 hours | 24 hours | 48 hours | 2 hours | 12 hours | 24 hours | 48 hours
| Avg | 198.715 | 491.4 | 949.725 | 1434.3 | 36.2475 | 91.8775 | 199.2 | 354.175
| SD | 230.7891 | 146.0567 | 507.5378 | 800.3576 | 21.4553 | 43.238 | 44.8379 | 233.9687
| ttest | 2 hours | 12 hours | 24 hours | 48 hours
| CSBv+tSpacerv+t v Spacerv+t | 0.21051 | 0.0019277 | 0.025745 | 0.041174
VANC
|  | CSBv+tSpacerv+t - 2 | CSBv+tSpacerv+t - 12 | CSBv+tSpacerv+t - 24 | CSBv+tSpacerv+t - 48 | Spacerv+t - 2 | Spacerv+t - 12 | Spacerv+t - 24 | Spacerv+t - 48 | VP -2 | VP - 12 | VP - 24 | VP - 48
| 1 | 164.7 | 1542 | 3272 | 5971 | 4.745 | 14.54 | 26.3 | 49.82 | 1470 | 1480 | 1492 | 1515
| 2 | 22.71 | 1904 | 3830 | 7038 | 21.62 | 86.53 | 113.1 | 113.1 | 1237 | 1247 | 1259 | 1282
| 3 | 314.5 | 742.8 | 1164 | 1164 | 33.88 | 70.04 | 85.91 | 85.91 | 15.83 | 25.63 | 31.51 | 60.91
| 4 | 788.6 | 1716 | 6222 | 20986 | 25.07 | 44.91 | 196.7 | 589.9
| Avg | 322.6275 | 1476.2 | 3622 | 8789.75 | 21.3287 | 54.005 | 105.5025 | 209.6825 | 907.61 | 917.5433 | 927.5033 | 952.6367
| SD | 332.7101 | 510.7909 | 2079.0645 | 8522.8039 | 12.2016 | 31.3858 | 70.7851 | 254.8 | 781.0416 | 781.1557 | 784.6498 | 780.9959
| ttest | 2 hours | 12 hours | 24 hours | 48 hours
| CSBv+tSpacerv+t v Spacerv+t | 0.12028 | 0.0014353 | 0.014843 | 0.090834
| CSBv+tSpacerv+t v VP | 0.22764 | 0.30005 | 0.090494 | 0.18184
| Spacerv+t v VP | 0.065665 | 0.071025 | 0.083648 | 0.12697
|  | Tobramycin | [7 empty] | Vancomycin
|  | 2 hours |  | 12 hours |  | 24 hours |  | 48 hours |  | 2 hours |  |  | 12 hours |  |  | 24 hours |  |  | 48 hours
|  | CSBv+tSpacerv+t | Spacerv+t | CSBv+tSpacerv+t | Spacerv+t | CSBv+tSpacerv+t | Spacerv+t | CSBv+tSpacerv+t | Spacerv+t | CSBv+tSpacerv+t | Spacerv+t | VP | CSBv+tSpacerv+t | Spacerv+t | VP | CSBv+tSpacerv+t | Spacerv+t | VP | CSBv+tSpacerv+t | Spacerv+t | VP
| AUC | 198.715 | 36.2475 | 491.4 | 91.8775 | 949.725 | 199.2 | 1434.3 | 354.175 | 322.6275 | 21.3287 | 907.61 | 1476.2 | 54.005 | 917.5433 | 3622 | 105.5025 | 927.5033 | 8789.75 | 209.6825 | 952.6367
| Std Error | 115.3945 | 10.7276 | 73.0283 | 21.619 | 253.7689 | 22.4189 | 400.1788 | 116.9843 | 166.355 | 6.1008 | 450.9346 | 255.3955 | 15.6929 | 451.0005 | 1039.5323 | 35.3926 | 453.0178 | 4261.4019 | 127.4 | 450.9082
LOG
|  | CSBv+tSpacerv+t - 2 | CSBv+tSpacerv+t - 12 | CSBv+tSpacerv+t - 24 | CSBv+tSpacerv+t - 48 | Spacerv+t - 2 | Spacerv+t - 12 | Spacerv+t - 24 | Spacerv+t - 48
| 1 | 1.5458 | 2.7052 | 3.1937 | 3.345 | 1.1824 | 1.817 | 2.3292 | 2.5966
| 2 | 1.4365 | 2.6991 | 2.9775 | 3.1804 | 1.7453 | 1.954 | 2.1222 | 2.1222
| 3 | 2.3263 | 2.4787 | 2.5041 | 2.5041 | 1.731 | 2.1861 | 2.3508 | 2.3508
| 4 | 2.7163 | 2.8176 | 2.986 | 3.2279 | 1.3077 | 1.7667 | 2.3553 | 2.8228
|  | 2 hours | 12 hours | 24 hours | 48 hours | 2 hours | 12 hours | 24 hours | 48 hours
| Avg | 2.0062 | 2.6752 | 2.9153 | 3.0643 | 1.4916 | 1.931 | 2.2894 | 2.4731
| SD | 0.61734 | 0.14185 | 0.29182 | 0.37986 | 0.28931 | 0.18761 | 0.11202 | 0.30308
| ttest | 2 hours | 12 hours | 24 hours | 48 hours
| CSBv+tSpacerv+t v Spacerv+t | 0.18187 | 0.00072803 | 0.0070794 | 0.050935
VANC
|  | CSBv+tSpacerv+t - 2 | CSBv+tSpacerv+t - 12 | CSBv+tSpacerv+t - 24 | CSBv+tSpacerv+t - 48 | Spacerv+t - 2 | Spacerv+t - 12 | Spacerv+t - 24 | Spacerv+t - 48 | VP -2 | VP - 12 | VP - 24 | VP - 48
| 1 | 2.2167 | 3.1881 | 3.5148 | 3.776 | 0.67624 | 1.1626 | 1.42 | 1.6974 | 3.1673 | 3.1703 | 3.1738 | 3.1804
| 2 | 1.3562 | 3.2797 | 3.5832 | 3.8474 | 1.3349 | 1.9372 | 2.0535 | 2.0535 | 3.0924 | 3.0959 | 3.1 | 3.1079
| 3 | 2.4976 | 2.8709 | 3.066 | 3.066 | 1.5299 | 1.8453 | 1.934 | 1.934 | 1.1995 | 1.4087 | 1.4984 | 1.7847
| 4 | 2.8969 | 3.2345 | 3.7939 | 4.3219 | 1.3992 | 1.6523 | 2.2938 | 2.7708
| Avg | 2.2418 | 3.1433 | 3.4895 | 3.7528 | 1.235 | 1.6494 | 1.9253 | 2.1139 | 2.4864 | 2.5583 | 2.5907 | 2.691
| SD | 0.65305 | 0.18542 | 0.30632 | 0.51806 | 0.38128 | 0.34555 | 0.36863 | 0.46223 | 1.1151 | 0.99623 | 0.94668 | 0.78572
| ttest | 2 hours | 12 hours | 24 hours | 48 hours
| CSBv+tSpacerv+t v Spacerv+t | 0.037386 | 0.00026646 | 0.00061741 | 0.0032542
| CSBv+tSpacerv+t v VP | 0.7273 | 0.28918 | 0.12724 | 0.081458
| Spacerv+t v VP | 0.085017 | 0.14263 | 0.24607 | 0.27217
|  | Tobramycin | [7 empty] | Vancomycin
|  | 2 hours |  | 12 hours |  | 24 hours |  | 48 hours |  | 2 hours |  |  | 12 hours |  |  | 24 hours |  |  | 48 hours
|  | CSBv+tSpacerv+t | Spacerv+t | CSBv+tSpacerv+t | Spacerv+t | CSBv+tSpacerv+t | Spacerv+t | CSBv+tSpacerv+t | Spacerv+t | CSBv+tSpacerv+t | Spacerv+t | VP | CSBv+tSpacerv+t | Spacerv+t | VP | CSBv+tSpacerv+t | Spacerv+t | VP | CSBv+tSpacerv+t | Spacerv+t | VP
| AUC | 2.0062 | 1.4916 | 2.6752 | 1.931 | 2.9153 | 2.2894 | 3.0643 | 2.4731 | 2.2418 | 1.235 | 2.4864 | 3.1433 | 1.6494 | 2.5583 | 3.4895 | 1.9253 | 2.5907 | 3.7528 | 2.1139 | 2.691
| Std Error | 0.30867 | 0.14466 | 0.070926 | 0.093803 | 0.14591 | 0.05601 | 0.18993 | 0.15154 | 0.32653 | 0.19064 | 0.64382 | 0.092709 | 0.17278 | 0.57517 | 0.15316 | 0.18432 | 0.54656 | 0.25903 | 0.23111 | 0.45364